## AX
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
Scan Head with Conversion Units for DUX- Choose 1
98721 | AX-R 2K SCAN HEAD & CONTROLLER | Scanhead,AX | AxR 2k Resonant Scanhead | AddPart:MHA52100,MHE60310
98722 | AX UPGRADABLE SCAN HEAD & CONTROLLER | Scanhead,AX | Resonant-Ready Ax Scanhead. 2K resonant can be added at a later date. | AddPart:MHA52000,MHE60320
98723 | AX GALVO-ONLY SCAN HEAD & CONTROLLER | Scanhead,AX | Ax Galvo-only Scanhead | AddPart:MHA52000,MHE60320
MHA52000 | AX-SHS AX SCAN HEAD & CONTROLLER | Scanhead,AX |  | AXScanHead:1:1:P | NSPARCGalvo:0:1:H | AXLaserIn1:1:1:H | AXLaserIn2:0:1:H | AXDualLaserCable:0:1:H | FirstDichroic:0:6:H | AXInterlockBox:1:1:H | AXConvUnit:0:1:H | AXClamp:0:2:H | PowerCord:1:1:H
MHA52100 | AX-SHR AX R SCAN HEAD & CONTROLLER | Scanhead,AX |  | AXScanHead:1:1:P | NSPARC:0:1:H | AXLaserIn1:1:1:H | AXLaserIn2:0:1:H | AXDualLaserCable:0:1:H | FirstDichroic:0:6:H | AXInterlockBox:1:1:H | AXRConvUnit:0:1:H | AXClamp:0:2:H | PowerCord:1:1:H
MHE60200 | DUX-ST DETECTOR UNIT |  |  | DUXPort:1:1:P | DUXFilterCube:1:18:H | DUXPMT:2:4:H
MHE60000 | DUX-VB4 DUVB 4CH DETECTOR UNIT |  |  | DUXPort:1:1:P | DUXFilterCube:1:12:H | DUXPMT:4:4:H
MHE60100 | DUX-VB2 DUVB 2CH DETECTOR UNIT |  |  | DUXPort:1:1:P | DUXFilterCube:1:6:H | DUXPMT:2:2:H
MHE60210 | PMT-MA MULTI ALKALI UNIT |  |  | DUXPMT:1:1:P
MHE60220 | PMT-GA GAASP UNIT |  |  | DUXPMT:1:1:P
MHE60230 | PMT-IRGA NIR GAASP DETECTOR |  |  | DUXPMT:1:1:P
MHE60240 | PMT-GAS GAAS DETECTOR |  |  | DUXPMT:1:1:P
98701 | DUX-ST 2CH MA |  | Fixed 2-channel detector with 2 High Sensitivty PMTs. Has one 6-position filterwheel | DUXPort:1:1:P | DUXFilterCube:1:6:H
98702 | DUX-VB 2CH GaAsP |  | Tunable 2-channel detector with 2 GaAsPs. Has a 6-position filterwheel in position 1 and one fully tunable detector in position 2. | DUXPort:1:1:P | DUXFilterCube:1:6:H
98703 | DUX-ST 4CH 2 GaAsP + 2MA |  | Fixed 4-channel detector with 2 High-Sensitivity PMTs in positions 1 and 4, 2 GaAsPs in positions 2 and 3. | DUXPort:1:1:P | DUXFilterCube:1:18:H
98704 | DUX-VB 4CH GaAsP |  | 4-Channel GaAsP detector with 6-position filterwheels in positions 1 and 4, and fully tunable detectors in positions 2 and 3. | DUXPort:1:1:P | DUXFilterCube:1:12:H
98706 | DUX-VB 4CH 3 GaAsP + 1NIR GaAsP |  | 4-Channel detector with 6-position filterwheels in positions 1 and 4, and fully tunable detectors in positions 2 and 3. Designed for use with NIR. Include one NIR-GaAsP in position 4 and GaAsP in postions 1-3. | DUXPort:1:1:P | DUXFilterCube:1:12:H
98707 | DUX-VB 4CH 3 GaAsP + 1GaAs |  | 4-Channel detector with 6-position filterwheels in positions 1 and 4, and fully tunable detectors in positions 2 and 3. Designed for use with NIR. Include one GaAs in position 4 and GaAsP in postions 1-3. | DUXPort:1:1:P | DUXFilterCube:1:12:H
98705 | DUX-VB 4CH 3 GaAsP + 1MA |  | 4-Channel detector with 6-position filterwheels in positions 1 and 4, and fully tunable detectors in positions 2 and 3. Designed for use with NIR (730). Include one high-sensitivity PMT in position 1 and GaAsP in postions 2-4. | DUXPort:1:1:P | DUXFilterCube:1:12:H
MHE60320 | AD-GAL CONVERSION UNIT |  | Quote ONLY for adding DUX to an AX-NSPARC standalone system. This product is included in 98722/98723. | AXConvUnit:1:1:P | DUXPort:1:1:H
MHE60310 | AD-R2K CONVERSION UNIT |  | Quote ONLY for adding DUX to an AXR-NSPARC standalone system. This product is included in 98721. | AXRConvUnit:1:1:P | DUXPort:1:1:H
NSPARC Alone Scanheads
98725 | AXR NSPARC SYSTEM | Scanhead,AX | Includes AXR scanhead + NSPARC. Must quote MHE60310 to add a DUX to this system. | AddPart:MHA52100,MHE62000
98726 | AX-GALVO NSPARC SYSTEM | Scanhead,AX | Includes AX scanhead + NSPARC. Must quote MHE60320 to add a DUX to this system. | AddPart:MHA52000,MHE62000,MXA22211
NSPARC + Adapters
MHE62000 | AX-ISM-VIS NSPARC | NSPARC,NIS-3D-DECON? | Includes quad barrier + DAPI, FITC, TRITC, CY5 single emitters. Quote this to add NSPARC to a DUX system. | NSPARC:1:1:P | NSPARCFilter:0:7:H | NSPARCUprightAdapter:0:1:H
MXA22211 | S-AX-ISM Galv.Connection Kit |  | Required to add NSPARC to Ax Galvo system (98723) | NSPARCGalvo:1:1:P | NSPARC:1:1:H
MHV55800 | AX-ISM-NI NIE ADAPTER | NSPARCUprightAdapter | Adapter required for upright configurations | NSPARCUprightAdapter:1:1:P
NSPARC Optional Filters
MHE63000 | AX-ISM-BA 482/35 BA FILTER |  |  | NSPARCFilter:1:1:P
MHE63010 | AX-ISM-BA 540/30 BA FILTER |  |  | NSPARCFilter:1:1:P
MHE63020 | AX-ISM-BA 607LP BA FILTER |  |  | NSPARCFilter:1:1:P
MHE63030 | AX-ISM-BA 523/610 DUAL FILTER |  |  | NSPARCFilter:1:1:P
77075042 | NSPARC CFP FILTER (470/24) |  |  | NSPARCFilter:1:1:P
77075043 | NSPARC GFP FILTER (510/20) |  |  | NSPARCFilter:1:1:P
77075044 | NSPARC YFP FILTER (540/24) |  |  | NSPARCFilter:1:1:P
77075045 | NSPARC RFP FILTER (595/44) |  |  | NSPARCFilter:1:1:P
77075038 | NSPARC RED FILTER (561LP) |  |  | NSPARCFilter:1:1:P
77075046 | NSPARC FAR RED FILTER (700/75) |  |  | NSPARCFilter:1:1:P
77075031 | NSPARC NIR FILTER (750LP) |  |  | NSPARCFilter:1:1:P
Laser Units - Choose 1
MHF49030 | LUD-S4 LASER UNIT 405/488/561/640 | LaserSafety |  | AXLaserIn1:1:1:P | LUDKit:1:1:H | LUD44Cable:0:1:H | InterlockPlug:1:1:P | PowerCord:1:1:H
MHF49050 | LUD-S6 LASER UNIT 405/445/488/514/561/64 | LaserSafety |  | AXLaserIn1:1:1:P | LUDKit:1:1:H | LUD62Cable:0:1:H | InterlockPlug:1:1:P | PowerCord:1:1:H
MHF46810 | LUD-AX AX CONNECTION KIT D |  |  | LUDKit:1:1:P
MHF49000 | LUA-S4 Laser Unit AS405/488/561/640 | LaserSafety |  | AXLaserIn1:1:1:P | LUAKit:1:1:H | LUA44Cable:0:1:H | InterlockPlug:1:1:P | PowerCord:1:1:H
MHF49020 | LUA-S6 LASER 405/445/488/514/561/640 | LaserSafety |  | AXLaserIn1:1:1:P | LUAKit:1:1:H | LUA62Cable:0:1:H | InterlockPlug:1:1:P | PowerCord:1:1:H
MHF46800 | LUA-AX AX Connection Kit A |  |  | LUAKit:1:1:P
LUNF Units for Ax
77097075 | LUNF LASER COMB 445/515/594/730 AX | LaserSafety |  | AXLUNF:1:1:P |  | LUNF44Cable:1:1:H | InterlockLUNF:1:1:H | PowerCord:1:1:H
77097058 | LUNF LASER 445/514 1F F/A1 SYSTEMS | LaserSafety |  | AXLUNF:1:1:P |  | LUNF44Cable¦LUNF62Cable:1:1:H | InterlockLUNF:1:1:H | PowerCord:1:1:H
77097078 | AX 730NM LASER SYSTEM | LaserSafety |  | AXLUNF:1:1:P |  | LUNF44Cable¦LUNF62Cable:1:1:H | InterlockLUNF:1:1:H | PowerCord:1:1:H
77097091 | LUNF 785NM 70MW 1F for AX | LaserSafety |  | AXLUNF:1:1:P |  | LUNF44Cable¦LUNF62Cable:1:1:H | InterlockLUNF:1:1:H | PowerCord:1:1:H
77012516 | AX LUN/LUA 4+4 LASER CONTROL CABLE |  | For use with LUN4 or LUAS4 + LUNF | AXDualLaserCable:1:1:P | LUA44Cable:1:1:P | LUNF44Cable:1:1:P
77012521 | AX LUA 6+2 LASER CONTROL CABLE |  | For use with LUA-S5 or LUA-S6 + LUNF | AXDualLaserCable:1:1:P | LUA62Cable:1:1:P | LUNF62Cable:1:1:P
77012528 | AX LUD-S4+4 LASER CONTROL CABL |  | For use with LUD-S and LUNF | AXDualLaserCable:1:1:P | LUD44Cable:1:1:P | LUNF44Cable:1:1:P
77012529 | AX LUD-S6+2 LASER CONTROL CABLE |  | For use with LUD-S and LUNF | AXDualLaserCable:1:1:P | LUD62Cable:1:1:P | LUNF62Cable:1:1:P
77023043 | AX PRISM 2ND INPUT 50/50 BEAMSPLITTER |  | Requires MHV56200 | AX5050:0:1:P
MHV56200 | AX-ILP AX OPTIONAL INPUT LASER PORT |  | *Includes 100/0 mirror | AXLaserIn2:0:1:P | AX5050:0:1:H | AXLUNF:1:1:H
MHV56500 | AX-ILP-IRDM AX LASER INPUT IR DM |  | Can be used in place of 77023043 + MHV56200 for 730 or 785 single line LUNF | AXLaserIn2:0:1:P |  | AXLUNF:1:1:H
1st Dichroic
98301 | AX FILTER 405/445/514/594/730 |  |  | FirstDichroic:1:1:P
98302 | AX FILTER 405/488/561/640/730 |  |  | FirstDichroic:1:1:P
MHE56120 | AX 1ST DICHROIC MIRROR 445 / 514 |  |  | FirstDichroic:1:1:P
MHE56110 | AX 1ST DICHORIC MIRROR 405/ 488 |  |  | FirstDichroic:1:1:P
MHE56140 | AX 1ST DM 405/488/561/640/785 |  |  | FirstDichroic:1:1:P
98303 | AX DUX ST2 V4 CUBE |  | Virtual 4-channel cube with fixed barriers for 405/640 in PMT1 and 488/561 in PMT2 | DUXFilterCube:1:1:P
98306 | AX DUX ST2 445/514 |  | Fixed 2-channel cube for CFP/YFP | DUXFilterCube:1:1:P
98307 | AX DUX ST2 488/561 |  | Fixed 2-channel cube for GFP/RFP | DUXFilterCube:1:1:P
98308 | AX DUX CFP 445 |  | Long pass filter for sending CFP emission to fixed detector position | DUXFilterCube:1:1:P
98309 | AX DUX ST4 YFP/MCHERRY514/594 |  |  | DUXFilterCube:1:1:P
98310 | AX DUX-VB V4 CUBE |  | Virtual 4-channel cube with fixed barrier for 405/640 in PMT1 and tunable passthrough for 488/561 | DUXFilterCube:1:1:P
98312 | AX DUX-VB 640SP |  | Short Pass Filter for sending 640-based emission to fixed detector position | DUXFilterCube:1:1:P
98313 | AX FILTER DUX-VB 730SP |  | Short Pass Filter for sending 730-based emission to fixed detector position | DUXFilterCube:1:1:P
99343 | A1/C2 447/57 EMISSION SET FOR 405 |  | Fixed DAPI cube | DUXFilterCube:1:1:P
99344 | A1/C2 521/42 EMISSION SET FOR 488 |  |  | DUXFilterCube:1:1:P
99345 | A1/C2 600/45+650LP FOR 561/640 EM |  | Fixed 2 channel cube for RFP/Far Red | DUXFilterCube:1:1:P
77015966 | FF01-785/62-25 EM FLTR MTD EM POS TI CBE |  | Emission filter for 730nm Excitation | DUXFilterCube:1:1:P
MHE57210 | AX-F 805LP FILTER CUBE |  | SP filter for use with 785nm Excitation | DUXFilterCube:1:1:P
Transmitted Detector Configuration
MHE50350 | AX-DUT DIASCOPIC DETECTOR UNIT |  | *Not compatible with nSPARC-only detector systems | DiaFilterSlider:0:1:P
MEE55855 | TI2-DUAX-LHLED AX LED LAMPHOUSE |  | Lamphouse REQUIRED for use with Transmitted Dectector on Ti2 | LamphouseE:1:1:P
MBF72210 | FN-LPA LAMPHOUSE ADAPTER |  | For FN1
Mounting Adapters and Interlock
MBD73100 | FN-S10 10MM SPACER FOR FN1 (SET OF 4) |  | Can be added for additional workspace between the objective and stage on an FN1
MHV55010 | A1 ADAPTER SET FOR TI |  | Required | SidePortAdapter:1:1:P | AXScanHead:1:1:H

## Ti2
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16 | slot17 | slot18 | slot19 | slot20 | slot21 | slot22 | slot23 | slot24 | slot25 | slot26 | slot27 | slot28 | slot29 | slot30
Stands
MEA54200 | TI2-U INVERTED MICROSCOPE MAIN BODY | LaserSafety,Ti2,Base |  | EyepieceTubeU:1:1:H | StageUpU:0:1:H | EpiTurretA:1:1:H | SidePortAdapter:0:2:H | Pillar:1:1:H | ACAdapter:1:1:H | FieldStop:0:1:H | NosePieceU:1:1:H | Analyzer:0:1:H | LAPPU:0:1:H | FixPlate:0:1:H | StageU¦StageH117:1:1:H |  |  | InterlockRound:0:1:H | CondenserU:1:1:H | AddPart:JBAM-Ti2SidePort,JBAM-Ti2SidePort
MEA54220 | TI2-U INVERTED E20L80 MAIN BODY | LaserSafety,Ti2,Base |  | Alias:MEA54200
MEA54100 | TI2-A INVERTED MICROSCOPE MAIN BODY | LaserSafety,Ti2,Base |  | EyepieceTubeA¦EyepieceTubeU:1:1:H | StageUpU:0:1:H | EpiTurretA:1:1:H | SidePortAdapter:0:2:H | Pillar:1:1:H | ACAdapter:1:1:H | FieldStop:0:1:H | NosePieceA¦NosePieceU:1:1:H | Analyzer:0:1:H | LAPPA¦LAPPU:0:1:H | FixPlate:0:1:H | StageU¦StageH117:1:1:H | AUX:0:6:H |  | InterlockRound:0:1:H | CondenserA¦CondenserU:1:1:H | AddPart:JBAM-Ti2SidePort,JBAM-Ti2SidePort
MEA54120 | TI2-A INVERTED E20L80 MAIN BODY | LaserSafety,Ti2,Base |  | Alias:MEA54100
MEA54000 | TI2-E INVERTED MICROSCOPE MAIN BODY | LaserSafety,Ti2,Base |  | EyepieceTubeE¦EyepieceTubeA¦EyepieceTubeU:1:1:H | StageUpE:0:1:H | EpiTurretE¦EpiTurretA:1:1:H | SidePortAdapter:0:2:H | Pillar:1:1:H | Ti2EController:1:1:H | FieldStop:0:1:H | NosePieceE¦NosePieceA¦NosePieceU:1:1:H | Analyzer:0:1:H | LAPPE¦LAPPA¦LAPPU:0:1:H | FixPlate:0:1:H | StageE¦StageU¦StageH117:1:1:H | AUX:0:6:H | JoystickE:1:1:H | InterlockRound:0:1:H | CondenserE¦CondenserA¦CondenserU:1:1:H | AddPart:JBAM-Ti2SidePort,JBAM-Ti2SidePort
MEA54010 | TI2-E/B INVERTED MICROSCOPE MAIN BODY | LaserSafety,Ti2,Base |  | Alias:MEA54000 | CmountAdapter:0:1:H
MEF55037 | TI2-CTRE TI2-E CONTROLLER |  | Required for Ti2-E Stands | Ti2EController:1:1:P | USBCord:1:1:H | PowerCord:1:1:H
Eyepiece Base Units; Bodytubes; Eyepieces
MEB55835 | TI2-T-BS S EYEPIECE TUBE BASE UNIT |  |  | EyepieceTubeU:1:1:P | BinoTube:1:1:H
MEB55830 | TI2-T-BS S EYEPIECE TUBE BASE UNIT |  |  | Alias:MEB55835
MEB55855 | TI2-T-BC PHOTO PORT TUBE BASE UNIT | PhotoPort |  | EyepieceTubeU:1:1:P | BinoTube:1:1:H | FmountAdapter:0:1:H
MEB55860 | TI2-T-BA ASSIST TUBE BASE UNIT |  | Ti2-A/E Only | EyepieceTubeA:1:1:P | BinoTube:1:1:H
MEB52340 | TC-T-TS S BINOCULAR TUBE |  | Requires Base Unit | BinoTube:1:1:P | EyePiece:2:2:H
MEB52520 | TC-T-ER ERGO BINOCULAR TUBE |  | Requires Base Unit | Alias:MEB52340
MAK10110 | CFI 10X EYEPIECE F.N. 22MM-NC |  |  | EyePiece:1:1:P | EyeCup:0:1:H
MAK99000 | CFI EYECUP F/22& 25MM (10MM HIGH) |  |  | EyeCup:1:1:P
MAK10120 | CFI 12.5X EYEPIECE |  |  | EyePiece:1:1:P | EyeCup:0:1:H
MAK30105 | CFI UW EYEPIECE 10X F.N. 25MM-NC |  |  | EyePiece:1:1:P | EyeCup:0:1:H
MAK85005 | C-CT CENTERING TELESCOPE-NC |  | Required for Phase or NAMC alignment on Ti2U | MISC:1:1:P
External Phase Unit (see condensers for compatibility)
MEB55840 | TI2-T-BP-E E-TUBE BASE UNIT EXTERNAL PH | XPhase,PhotoPort |  | EyepieceTubeE:1:1:P | BinoTube:1:1:H | FmountAdapter:0:1:H | ExternalPhaseSULens:0:1:H | ExternalPhaseModule:1:5:H
MEB55845 | TI2-T-BP-E E-TUBE BASE UNIT EXTERNAL PH | XPhase,PhotoPort |  | Alias:MEB55840
MEH51695 | TI2-T-XPH-PH3 40X PH3 RING |  |  | ExternalPhaseModule:1:1:P | PH3Module:1:1:H
MEH51700 | TI2-T-XPH-PH3 60X PH3 RING |  |  | ExternalPhaseModule:1:1:P | PH3Module:1:1:H
MEH51710 | TI2-T-XPH-PH4 60X PH4 RING |  |  | ExternalPhaseModule:1:1:P | PH4Module:1:1:H
MEH51720 | TI2-T-XPH-PH3 100X PH3 RING |  |  | ExternalPhaseModule:1:1:P | PH3Module:1:1:H
MEH51730 | TI2-T-XPH-PH4 100X PH4 RING |  |  | ExternalPhaseModule:1:1:P | PH4Module:1:1:H
MEV54005 | TI2-T-SULXPH STAGE UP LENS EXTERNAL PH | XPhaseSUL |  | ExternalPhaseSULens:1:1:P
Nosepieces
MEP51316 | TI2-N-N SEXTUPLE NOSEPIECE |  |  | NosePieceU:1:1:P | CFI60Objective:1:6:H | NosepieceTray:0:1:H
MEP53320 | TI2-N-ND DIC SEXTUPLE NOSEPIECE |  |  | NosePieceU:1:1:P | CFI60Objective:1:6:H | NosepieceTray:0:1:H
MEP53325 | TI2-N-ND-I INTELL DIC SEXTUPLE NOSEPIECE |  | Ti2 A or E Only | NosePieceA:1:1:P | CFI60Objective:1:6:H | NosepieceTray:0:1:H
MXA22161 | TI2-N-PT NOSEPIECE PROTECTION TRAY |  |  | NosepieceTray:1:1:P
MEP59320 | TI2-N-ND-E MOTOR DIC SEXTUPLE NOSEPIECE |  | Ti2 E Only | NosePieceE:1:1:P | CFI60Objective:1:6:H
Perfect Focus Nosepiece
MEP59394 | TI2-N-ND-P PERFECT FOCUS MOTOR NOSEPIECE |  | Ti2-E Only | NosePieceE:1:1:P | CFI60Objective:1:6:H | WID:0:1:H |  | PFSController:0:1:H
MEP59395 | TI2-N-NDM-P MP PERFECT FOCUS MTR NP MP |  | Ti2-E Only | NosePieceE:1:1:P | CFI60Objective:1:6:H | WID:0:1:H |  | PFSController:0:1:H
MEP59396 | TI2-N-NDA-P PFS MOTOR NOSEPIECE ACC |  | Ti2-E Only | NosePieceE:1:1:P | CFI60Objective:1:4:H | WID:0:1:H | CFI60ACCObjective:1:1:H | PFSController:0:1:H
MXA22190 | TI2-PFS-REMOTE OFFSET CONTROLLER |  | Optional | PFSController:1:1:P
MEV54006 | TI2-N-WID WATER IMMERSION DISPENSER |  | Ti2-E Only. Must be used in conjunction with PFS Nosepiece, Nikon Motorized Stage and Water Immersion Objective. | WID:2:2:P
Diascopic Components
MEE59925 | TI2-D-PD DIASCOPIC ILLUMINATION PILLAR |  |  | Pillar:1:1:P | LamphouseU¦LamphouseE:1:1:H | DiaFilterSlider:0:1:H | ContrastLightShield:0:1:H | DiaFilter:0:1:H
MEE59920 | TI2-D-PD DIASCOPIC ILLUMINATION PILLAR |  |  | Alias:MEE59925
MEE55705 | TI2-D-LHLED TRANSMITTED LED LAMP HOUSE |  | NOT compatible with confocal transmitted detectors | LamphouseU:1:1:P | SHADI:0:1:H
MEE55700 | TI2-D-LHLED TRANSMITTED LED LAMP HOUSE |  |  | Alias:MEE55705
MEE55800 | TI2-DUCF-LHLED LED LAMP HOUSE FOR CF DIA |  |  | LamphouseE:1:1:P
MEE55805 | TI2-DUCF-LHLED LED LAMP HOUSE FOR CF DIA |  | Preferred Lamphouse for use with Transmitted Dectector on Ti2 and Confocal | LamphouseE:1:1:P
MBE75225 | D-LH/LC PRECENTERED LAMPHOUSE W/ LC-NC |  |  | LamphouseU:1:1:P
MEV50015 | TI2-D-D 45MM DIFFUSER FILTER |  | Recommended when Halogen Lamp House is used. | DiaFilter:1:1:P
MEV50010 | TI2-D-SF DIASCOPIC ILLUMINATION SLIDER |  |  | DiaFilterSlider:1:1:P | DiaFilter:1:2:H
MBN11710 | FILTER 45MM NCB11 |  | Recommended when Halogen Lamp House is used. | DiaFilter:1:1:P
MBN11500 | FILTER 45MM HEAT ABSORBING |  |  | DiaFilter:1:1:P
MBN11200 | FILTER 45MM GIF |  |  | DiaFilter:1:1:P
MBN21802 | FILTER 45MM ND2 A |  |  | DiaFilter:1:1:P
MBN21804 | FILTER 45MM ND4 A |  |  | DiaFilter:1:1:P
MBN21808 | FILTER 45MM ND8 A |  |  | DiaFilter:1:1:P
MBN21816 | FILTER 45MM ND16 A |  |  | DiaFilter:1:1:P
Condensers & Phase Contrast
MEL47505 | TI2-C ELWD PHASE TURRET CONDENSER |  |  | CondenserU:1:1:P | AddPart:MEL37405,MEH52050,MEH51400,MEH51500 | CondenserModule:0:3:H
MEL51005 | TC-C-TC CONDENSER TURRET |  |  | CondenserU¦Ts2RCondenserAdapter:1:1:P | CondenserModule:0:7:H | CondenserLens¦CLWDA:1:1:H | PolarizerU:0:1:H
MEL51010 | TI2-C-TC-I INTELLIGENT CONDENSER TURRET |  | Ti2-A or E Only | CondenserA:1:1:P | CondenserModule:0:7:H | CondenserLens¦CLWDA:1:1:H | PolarizerA¦PolarizerU:0:1:H
MEL51920 | TI2-C-TC-E MOTORIZED CONDENSER TURRET |  | Ti2-E only | CondenserE:1:1:P | CondenserModule:0:7:H | CondenserLens¦CLWDA:1:1:H | PolarizerA¦PolarizerU:0:1:H
MEL37405 | TC-C ELWD CONDENSER LENS |  |  | CondenserLens:1:1:P
MEH52050 | TC-C-ME-PHL ELWD MODULE |  |  | CondenserModule:1:1:P | PHLModule:0:6:P | CondenserLens:1:1:H
MEH51400 | TC-C-ME-PH1 ELWD MODULE |  |  | CondenserModule:1:1:P | PH1Module:0:6:P
MEH51500 | TC-C-ME-PH2 ELWD MODULE |  |  | CondenserModule:1:1:P | PH2Module:0:6:P
MEL56200 | TI-C-LWD LWD LENS UNIT F/SYS COND TURRET |  |  | CondenserLens:1:1:P
MEH51050 | TC-C-ML-PHL LWD MODULE |  |  | CondenserModule:1:1:P | PHLModule:0:6:P
MEH51100 | TC-C-ML-PH1 LWD MODULE |  |  | CondenserModule:1:1:P | PH1Module:0:6:P
MEH51200 | TC-C-ML-PH2 LWD MODULE |  |  | CondenserModule:1:1:P | PH2Module:0:6:P
MEH51300 | TC-C-ML-PH3 LWD MODULE |  |  | CondenserModule:1:1:P | PH3Module:0:6:P
MXA22176 | TI2-C-CLWDA CLWD CONDENSER ADAPTER |  | Required for CLWD+Systems Condenser | CLWDA:1:1:P | CLWDLens:1:1:P
MEL56100 | TI-C-CLWD CLWD LENS F/SYS COND TURRET |  | Compatible with LWD N2 prism | CLWDLens:1:1:P
MEH51660 | TI2-C-MC-PH1 CLWD MODULE |  |  | CondenserModule:1:1:P | PH1Module:0:6:P
MEH51670 | TI2-C-MC-PH2 CLWD MODULE |  |  | CondenserModule:1:1:P | PH2Module:0:6:P
MEH51680 | TI2-C-MC-PH3 CLWD MODULE |  | Compatible with Full Intensity Phase | CondenserModule:1:1:P | PH3Module:0:6:P
MEH51690 | TI2-C-MC-PH4 CLWD MODULE |  | Compatible with Full Intensity Phase | CondenserModule:1:1:P | PH4Module:0:6:P
MEL59500 | T-CHNA HIGH NA COMMON COND LENS & EXTNDR |  |  | CondenserLens:1:1:P
MEL50020 | TI2-C-SCH HNA CONDENSER SLIDER |  |  | CondenserU:1:1:P
MEL51220 | TI2-C-LHD HNA DRY CONDENSER LENS |  |  | CondenserLens:1:1:P
MEL51420 | TI2-C-LHO HNA OIL CONDENSER LENS |  |  | CondenserLens:1:1:P
MXA22091 | TI-DF DARK-FIELD CONDENSER ADAPTER |  | Required for DF | CondenserU:1:1:P | DarkFieldCondenser:1:1:H
Nikon Adv. Modulation Contrast (NAMC)
MEL59910 | TI-C NAMC CONDENSER LENS |  |  | CondenserLens:1:1:P
MEN51943 | TC2-C-DICPNI NAMC/IMSI POLARIZER |  |  | PolarizerU:1:1:P
DIC & Pol Components
MEN51942 | TC-C-DICP DIC POLARIZER | Polarizer |  | PolarizerU:1:1:P
MEN51944 | TI2-C-DICP-I INTELLIGENT POLARIZER | Polarizer | Ti2-A or E Only | PolarizerA:1:1:P
MEN51928 | TI2-C-DICA DIC ANALYZER | Analyzer |  | Analyzer:1:1:P
MEH51510 | TC-C-ML-N1D LWD DRY DIC MODULE |  |  | CondenserModule:1:1:P | DICPrismN1:0:6:P
MEH51520 | TC-C-ML-N2D LWD DRY DIC MODULE |  | Compatible with CLWD Condenser System | CondenserModule:1:1:P | DICPrismN2:0:6:P
MEH51530 | TC-C-ML-NRD LWD DRY DIC MODULE |  |  | CondenserModule:1:1:P | DICPrismNR:0:6:P
MEH51540 | TI2-C-MH-N2D HNA DRY DIC MODULE |  |  | CondenserModule:1:1:P | DICPrismN2:0:6:P
MEH51550 | TI2-C-MH-NRD HNA DRY DIC MODULE |  | Compatible with CLWD Condenser System | CondenserModule:1:1:P | DICPrismNR:0:6:P
MEH51560 | TI2-C-MH-N2O HNA OIL DIC MODULE |  |  | CondenserModule:1:1:P | DICPrismN2:0:6:P
MEH51570 | TI2-C-MH-NRO HNA OIL DIC MODULE |  |  | CondenserModule:1:1:P | DICPrismNR:0:6:P
Episcopic Components
MEV51020 | TI2-F-FLT-I INTELLIGENT EPI FLUOR TURRET |  |  | EpiTurretA:1:1:P | FilterCube¦FilterCubeLFOV:0:6:H | DaisyPlugIn:1:1:H | DaisyPlugOut:0:1:H
MEV51030 | TI2-F-FLT-E MOTORIZED EPI FLUOR TURRET |  |  | EpiTurretE:1:1:P | FilterCube¦FilterCubeLFOV:0:6:H | DaisyPlugIn:1:1:H | DaisyPlugOut:0:1:H
MEV51040 | TI2-F-FLTH-E E-HQ MOTO HYQ EPI FL TURRET |  |  | Alias:MEV51030
MEE54405 | TI2-P-FWB-E BARRIER FILTER WHEEL |  |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter¦32mmFilter:0:7:H
MEE54410 | TI2-P-FWBS-E BARRIER FILTER WHEEL STGEUP |  |  | EpiTurretE:1:1:P | 25mmFilter¦32mmFilter:0:7:H
MXA22147 | S-TI2-DCL DAISY CABLE LONG |  |  | AUX:1:1:P | DaisyPlugIn:1:1:P
MXA22146 | S-TI2-DCS DAISY CABLE SHORT |  | Required for 2 turret stage up communication Ti2-A or E Only | DaisyPlugIn:1:1:P | DaisyPlugOut:1:1:P
Episcopic Illuminators
MEV50020 | TI2-F-FSC CIRCLE FIELD STOP SLIDER |  | Field of View Reduced to 22mm when in Optical Path | FieldStop:0:1:P
MEV50030 | TI2-F-FSS SQUARE FIELD STOP SLIDER |  |  | FieldStop:0:1:P
MEV50040 | TI2-F-FSR RECTANGULAR FIELD STOP SLIDER |  |  | FieldStop:0:1:P
D-FLEDI Nikon LED Fluorescence Illumination System
MBF83000 | D-LEDI FLUORESCENCE LED ILLUMINATION SYS |  |  | DLED:1:1:P | ACAdapter:1:1:H | DLEDTriggerCable:0:1:H
MBF83200 | D-LEDI2 FLUOR LED ILLUMINATION SYSTEM |  |  | DLED:1:1:P | ACAdapter:1:1:H | DLEDTriggerCable:0:1:H
MXA22196 | D-LEDI TI2 CONNECTION CABLE
77012522 | TRIGGER CABLE/D-LEDI W/BNC CONNECTORS |  | Requires 4 available TTL or AO as TTL on any nidaq device. | DLEDTriggerCable:1:1:P
Accessories for Laser Applications
MXA22152 | TI2-D-LSK LASER SAFETY KIT
Stages
MEC54100 | TI-SP PLAIN STAGE |  |  | StageU:1:1:P | StageRing:0:1:H
MXA22058 | GLASS STAGE RING - CENTER HOLE 32MM |  |  | StageRing:0:1:P
MXA20746 | TE ACRYLIC STAGE RING |  |  | StageRing:0:1:P
84450 | TPI-108RX THERMO PLATE F/TI,TS100 & TS2 |  | Tokai Hit | StageRing:0:1:P
MEC51200 | TC-S-SR RECTANGULAR MECHANICAL STAGE |  |  | StageU:1:1:P | StageUInsert:1:6:H | 96WellClamper:0:1:H | StageUHandle:1:1:H
MEC51210 | TC-S-SRF RECTANGULAR MECH STAGE -FRONT R |  |  | Alias:MEC51200
MEC57100 | TI2-S-HL LONG HANDLE |  |  | StageUHandle:1:1:P
MEC57110 | TC-S-HM MIDDLE HANDLE |  |  | StageUHandle:1:1:P
MEC57120 | TC-S-HS SHORT HANDLE |  |  | StageUHandle:1:1:P
MXA22142 | C-S-WC 96 WELL PLATE CLAMP |  |  | 96WellClamper:0:1:P
MEC55012 | C-S-HT TERASAKI PLATE HOLDER |  |  | StageUInsert:0:1:P
MEC55014 | C-S-HS SLIDE GLASS HOLDER |  |  | StageUInsert:0:1:P
MEC55016 | C-S-HU UNIVERSAL HOLDER |  |  | StageUInsert:0:1:P
MEC55017 | C-S-HP35 PETRI DISH HOLDER 35MM |  |  | StageUInsert:0:1:P
MEC55019 | TC-S-HA ACRYLIC HOLDER |  |  | StageUInsert:0:1:P
MEC55020 | C-S-HLS RING HOLDER SET |  |  | StageUInsert:0:1:P
MEC55021 | C-S-HLP100 PETRI DISH HOLDER 100MM |  |  | StageUInsert:0:1:P
MXA22143 | TC-S-SC STAGE CLIP |  |  | StageUInsert:0:1:P
84448 | TPI-TS2X THERMO PLATE F/TS2 MECH STAGE |  | Tokai Hit - Replaces stage plate insert holder of attachable mechanical stage | StageUInsert:0:1:P
84449 | TPI-TCSX THERMO PLATE F/TS2R & TI2 STAGE |  | Tokai Hit | StageUInsert:0:1:P
MEC56120 | TI2-S-SE-E MOTORIZED STAGE WITH ENCODER |  | No Controller Needed, Connects to TI2-E Control Box, Joystick Required Ti2-E Only | StageE:1:1:P | StageEInsert:1:6:H |  | WID:0:1:H
MEC56130 | TI2-S-SS-E MOTORIZED STAGE |  |  | Alias:MEC56120
MEF55706 | TI2-S-JS-SS JOYSTICK |  | Ti2-E Only | JoystickE:1:1:P
MEC59140 | TI2-S-HU UNIVERSAL HOLDER |  |  | StageEInsert:0:1:P
MEC59150 | TI2-S-HW WELL PLATE HOLDER |  |  | StageEInsert:0:1:P
90306 | SIX SLIDE HOLDER FOR Ti2 MOTORIZED STAGE |  |  | StageEInsert:0:1:P
MEC59160 | TI2-S-HJ STAGE RING HOLDER |  |  | StageEInsert:0:1:P | StageRing:0:1:H
90303 | 60MM PETRI DISH HOLDER F/MECH STAGE |  |  | StageRing:0:1:P
90304 | 35MM PETRI DISH HOLDER F/MECH STAGE |  |  | StageRing:0:1:P
90305 | 100MM PETRI DISH HOLDER F/MECH STAGE |  |  | StageRing:0:1:P
MXB22586 | TI2-S-HP POD HOLDER |  | For use with Emulate system only
90924 | TID-NA MOTORIZED STAGE ADAPTER
77033011 | ROTATING STAGE FOR TE300/2000/TI/TS2R |  | Rotating Gliding Stage. Movement is approximately 8.5mm from center in all dirrections
Expansion & Accessories
MBE96910 | NI-SHADI MTR SHUTTER ADPTR F/DIA ILLUM |  |  | SHADI:1:1:P | NISHE:1:1:H
MBE96000 | NI-SH-E MOTORIZED SHUTTER |  |  | NISHE:1:1:P | NISHCON:0:1:H
MBE96100 | NI-SH-CON CONTROLLER F/MOTORIZED SHUTTER |  |  | NISHE:1:1:P
MBE96905 | TI-SHAEP-I MTR SHUTTER ADPTR F/IN EPI-FL
MBE96930 | NI-SHCL MOTORIZED SHUTTER CABLE LONG
MXA22172 | S-TI2-SHCLL MOTOTRIZED SHUTTER CABLE LL |  | Connect to Ti2-E Control box for Trigger Control
MEF55806 | TI2-EB TI2-E EXTENSION BOX |  | Ti2-E Only
MED54550 | TI2-FP FIXED PLATE | FixingPlate |  | FixPlate:1:1:P
MED54555 | TI2-FP FIXED PLATE | FixingPlate |  | Alias:MED54550
77011387 | TI ENCODER ADAPTER KIT FOR NIKRFK
92080 | ECLIPSE MICROSCOPE PAD |  |  | MISC:0:1:P
Piezos
MXA22090 | TI PZT CONTROL CABLE

## Ts2
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
MFA34100 | ECLIPSE TS2 MAIN BODY | Base | Cannot be upgraded to Epi-Fl | Ts2ContrastSlider:0:1:H | EyePiece:2:2:H | Ts2CameraPort:0:1:H | Ts2EmbossSlider:0:1:H | CFI60Objective:1:4:H | StageRing:0:1:H | Ts2Stage:0:1:H | PowerCord:1:1:H
MFA34200 | ECLIPSE TS2-FL MAIN BODY | Base | Epi-FL Capability must be selected at time of purchase. | Alias:MFA34100 | Ts2LEDUnit:0:3:H | Ts2LightShield:0:1:H
MMK30105 | TS2-W 10x EYEPIECE |  |  | EyePiece:1:1:P
MFB36100 | TS2-P-CF CAMERA PORT 100 |  |  | Ts2CameraPort:1:1:P | CmountAdapter:1:1:H
MQD43035 | F-1X-TS2 TS2 F-MOUNT ADAPTER 1X |  |  | CmountAdapter:1:1:P | FmountAdapter:1:1:H
MQD43030 | C-0.63X-TS2 C MOUNT ADAPTER |  | Fits directly into MFB36100 TS2-P-CF Camera Port 100 | CmountAdapter:1:1:P | CmountCamera:1:1:H
MFV30010 | T1-SNCP PRECENT PH SLDR (PHL PH1 BF)-NC |  |  | Ts2ContrastSlider:1:1:P | SPH2:0:1:H | PHLModule:0:6:P | PH1Module:0:6:P
MFV30011 | T1-SCP CENTERABLE PH SLR (PHL,PH1,BF)-NC |  |  | Ts2ContrastSlider:1:1:P |  | PHLModule:0:6:P | PH1Module:0:6:P
MFV30013 | T1-SPH2 PHASE RING 2-NC |  |  | SPH2:1:1:P | PH2Module:0:6:P
MFH30010 | TS2-C-S-EC EMBOSS CONTRAST SLIDER |  | Objectives supported by Emboss Contrast are: MRP00102 CFI Achro Flat Field 10x, MRP00202 CFI Achro Flat Field LWD 20x and MRP05422 CFI Achro LWD 40xC | Ts2ContrastSlider:1:1:P
MFH30020 | TS2-C-ST-EC EMBOSS CONTRAST SLIDER |  | Required for Emboss Contrast | Ts2EmbossSlider:1:1:P
MBF84715 | C-LEDLH 385nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MBF84725 | C-LEDLH 455nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MBF84735 | C-LEDLH 470nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MBF84745 | C-LEDLH 505nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MBF84755 | C-LEDLH 525nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MBF84785 | C-LEDLH 560nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MBF84765 | C-LEDLH 590nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MBF84775 | C-LEDLH 625nm LED UNIT |  |  | Ts2LEDUnit:1:1:P | FilterCube:1:1:H
MFV51010 | TS2-LS LIGHT SHIELD |  |  | Ts2LightShield:1:1:P
MFV51000 | TS2-LS LIGHT SHIELD ELWD |  |  | Ts2LightShield:1:1:P
MFC10010 | TS2-S-SA AUXILIARY STAGE |  |  | Ts2Stage:1:1:P
MFC35004 | TS2-S-SM MECHANICAL STAGE |  |  | Ts2Stage:1:1:P | StageUInsert:1:10:H
MXA22066 | COVER TYPE 120 - TS/E600 POL-FN |  |  | MISC:1:1:P
MFA51000 | ECLIPSE TS2R MAIN BODY |  | Cannot be upgraded to Epi-Fl | PowerCord:1:1:H | PolarizerU:0:1:H | Ts2RCondenserAdapter:1:1:H | BinoTube:1:1:H | Ts2RCameraPort:0:1:H | Ts2REmbossSlider:0:1:H | Ts2RAnalyzer:0:1:H | CFI60Objective:1:6:H | StageU:1:1:H
MFA51010 | ECLIPSE TS2R-FL MAIN BODY |  | Epi-FL Capability must be selected at time of purchase. | Alias:MFA51000 | Ts2LightShield:0:1:H | Ts2LEDUnit:1:4:H | Ts2RLightAdapter:1:1:H
MFL51020 | TS2R-C-AN NAMC CONDENSER ADAPTER |  |  | Ts2RCondenserAdapter:1:1:P | CondenserU:1:1:H
MFL51025 | TS2R-C-AL LWD CONDENSER ADAPTER |  |  | Ts2RCondenserAdapter:1:1:P | CondenserU:1:1:H
MFL51015 | TS2R-C-SS SLIDER SLOT |  | Requires - MFV50010 or MVF50020 for Phase Contrast | Ts2RCondenserAdapter:1:1:P | CondenserLens:1:1:H | Ts2RPhaseSlider:0:1:H
MFV50010 | TS2R-C-SFPH NON-CENTERABLE PHASE SLIDER |  |  | Ts2RPhaseSlider:1:1:P | TS2R-C-MPH:0:1:H
MFV50030 | TS2R-C-MPH PH2 MODULE |  |  | TS2R-C-MPH:1:1:P
MFB56200 | TS2R-P-CF CAMERA PORT 100 |  |  | Ts2RCameraPort:1:1:P | CmountAdapter¦FmountAdapter:1:1:H
MFB56210 | TS2R-P-CH CAMERA PORT 80 |  |  | Ts2RCameraPort:1:1:P | FmountAdapter:1:1:H
MFB56220 | TS2R-P-CHL CAMERA PORT 50L |  |  | Ts2RCameraPort:1:1:P | FmountAdapter:1:1:H
MFH50020 | TS2R-C-SE-EC EMBOSS CONTRAST ELWD SLIDER |  | Required for Emboss w/ELWD | Ts2REmbossSlider:1:1:P
MFH50021 | TS2R-C-SL-EC EMBOSS CONTRAST LWD SLIDER |  | Required for Emboss w/LWD | Ts2REmbossSlider:1:1:P
MFN51929 | TS2R-C-DICA ANALYZER |  |  | Ts2RAnalyzer:1:1:P
MXA22164 | TS2R-C-SO SPINDLE OBSERVATION SYSTEM |  |  | Ts2RAnalyzer:1:1:P
MBF72320 | TS2R-F-AFL FIBER ADAPTER |  |  | Ts2RLightAdapter:1:1:P | LLG:1:1:H
MBF72330 | TS2R-F-AFL-2 D-LEDI ADAPTER |  |  | Ts2RLightAdapter:1:1:P | FIBA¦DLED:1:1:H

## Ni
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
MBA91000 | ECLIPSE NI-E MAIN BODY | LaserSafety,Upright,Base | Requires MBF95400 Control Box A, Arm & Substage | NIECAM¦NISAM:1:1:H | NISSR¦NISSF:1:1:H | NIEStageCompatibility:0:1:H | NIND2:0:1:H | NiBasePolarizer:0:1:H | LamphouseU:1:1:H | NIECB:0:1:H | InterlockRound:0:1:H |  |  | NIAMSCompatibility:0:1:H | NIECondenserCompatibility:0:1:P
MBA92020 | ECLIPSE NI-L MAIN BODY | Upright,Base |  | NIUCAM¦NISAM:1:1:H | NISSR¦NISS:1:1:H | NILStageCompatibility:0:1:H | NLSC:0:1:H | NiBasePolarizer:0:1:H |  | NICTLB:0:1:H |  | PowerCord:1:1:H
MBF95400 | NI-CTLA CONTROL BOX A |  | Required. Use with Halogen Lamphouse | NICTLA:1:1:P
MBF95450 | NI-CTLA2 CONTROL BOX A2 |  | Required. Use with High Color Rendering LED Lamphouse | NICTLA2:1:1:P
MBD92100 | NIE-CAM CONTACT ARM | NiLayer |  | NIECAM:1:1:P | NiNosepieceE¦NIAMS:1:1:H | UprightTube¦NITT¦NITTE¦NiIntermediateTube¦NIFLT6¦NIFLT6I¦NIFLT6E:1:1:H
MBD92000 | NI-SAM STANDARD ARM | NiLayer |  | NISAM:1:1:P | NiNosepieceU¦NIAMS:1:1:H | UprightTube¦NITT¦NiIntermediateTube¦NIFLT6:1:1:H
MBD92110 | NIU-CAM CONTACT ARM |  |  | NIUCAM:1:1:P | NiNosepieceE:1:1:H | UprightTube¦NITT¦NiIntermediateTube¦NIFLT6¦NIFLT6I¦NIFLT6E:1:1:H
MBP99405 | NI-N7 INTELLIGENT SEPTUPLE NOSEPIECE-NC |  | Requires MBD92100 Contact Arm | NiNosepieceE:1:1:P | CFI60Objective:1:7:H | NiAnalyzer:0:1:H
MBP98405 | NI-N7-E MOTORIZED SEPTUPLE NOSEPIECE-NC |  | Requires MBD92100 Contact Arm | NiNosepieceE:1:1:P | CFI60Objective:1:7:H | NiAnalyzer:0:1:H
MBP98325 | NI-ND6-E MOTOR DIC SEXTUPLE NOSEPIECE-CN |  | Requires MBD92100 Contact Arm | NiNosepieceE:1:1:P | CFI60Objective:1:6:H | NiAnalyzer:0:1:H | DLP:0:1:H
MBP71316 | C-N6 SEXTUPLE NOSEPIECE ESD-NC |  | use with MBD92000 Standard Arm | NiNosepieceU:1:1:P | CFI60Objective:1:6:H
MBP71325 | C-NA SEXTUPLE NOSEPCE W/ANALYZER SLOT-NC |  | use with MBD92000 Standard Arm | NiNosepieceU:1:1:P | CFI60Objective:1:6:H | NiAnalyzer:0:1:H
MBP73325 | D-ND6 DIC SEXTUPLE NOSEPIECE-NC |  | use with MBD92000 Standard Arm | NiNosepieceU:1:1:P | CFI60Objective:1:6:H | NiAnalyzer:0:1:H | DLP:0:1:H
MBP60125 | LV-NBD5BD-CH MANUAL QUIN NOSEPIECE ESD |  | use with MBD92000 Standard Arm | NiNosepieceU:1:1:P | CFI60Objective:1:5:H
MBP60115 | LV-NU5-CH UNIVER MAN QUIN NPC ESD F/DIC |  | use with MBD92000 Standard Arm | NiNosepieceU:1:1:P | CFI60Objective:1:5:H
MBP70100 | FN-S2N CFI60 TWO PLACE SLIDE NOSEPIECE |  |  | FNNosepiece:1:1:P | CFI60Objective:1:2:H
MBP70110 | FN-MN-H SINGLE CFI75 OBJECTIVE NOSEPIECE |  |  | FNNosepiece:1:1:P | CFI75Objective:1:1:H
MBP70120 | FN-MN-H2 SINGLE CFI90 OBJECTIVE NOSEPIECE |  |  | FNNosepiece:1:1:P | CFI90Objective:1:1:H
MBP70400 | FN-S2N-2 CFI90 TWO PLACE SLIDE NOSEPIECE |  |  | FNNosepiece:1:1:P | CFI90Objective:1:2:H
MBN74925 | D-DA SIMPLE ANALYZER |  | Installs in Nosepiece Slot | NiAnalyzer:1:1:P
MBN74920 | D-DA DIC ANALYZER |  | Installs in Nosepiece Slot | Alias:MBN74925
MBN91920 | NI-FA FL/DIC ANALYZER |  | For use with MBE94500 NI-FLT6 EPI-FLUORESCENCE CUBE TURRET Only | NIFAFL:1:1:P
MXA22120 | NI-AST ANALYZER SLIDER STOPPER |  |  | NIAST:1:1:P
MBN71924 | D-FA FL/DIC ANALYZER |  | For Use with MBE74100 D-FL UNIVERSAL EPI-FLUOR ILLUMINATOR Only | DFL2Slider:1:1:P
MBN74960 | D-LP 1/4 LAMBDA PLATE |  |  | DLP:1:1:P
MBF95420 | NIE-CB CONNECTOR BOX |  |  | NIECB:1:1:P | NISJ¦NIERG:0:1:H | NINDE:0:1:H | PowerCord:1:1:H
MBF95410 | NI-CTLB CONTROL BOX B |  |  | NICTLB:1:1:P | NISRCP:1:1:H
MBF95120 | NI-SRCP SIMPLE REMOTE CONTROL PAD |  |  | NISRCP:1:1:P
MBF95130 | NI-SJ JOYSTICK FOR MOTORIZED STAGE |  | Requires MBF95420 connector box C | NISJ:1:1:P
MBF95110 | NI-ERG NI ERGO CONTROLLER |  | Requires MBF95420 connector box C | NIERG:1:1:P
MBN98800 | NI-ND-E MOTORIZED ND FILTER UNIT |  | Requires MBF95420 connector box C | NINDE:1:1:P
MBN91802 | NI-ND2 FILTER 36MM ND2 |  |  | NIND2:1:1:P
MBE92020 | NL-SC SHUTTER CASSETTE FOR NI-L |  | Recommended for use when adding Epi-fluourescence to Ni-L stand | NLSC:1:1:P
MBE94610 | NI-EXW-E MOTORZD EXCITATION FILTER WHEEL |  | Requires MBF95420 Ni-E Aux Connector Box
MBE94620 | NI-BAW-E MOTORIZED BARRIER FILTER WHEEL |  | Fits underneath body tube
MBD92200 | NI-FNA FOCUSING NOSEPIECE SUB ARM |  | Requires either 49 or 74mm spacer, nosepiece, and MBD93030 substage | NIFNA:1:1:P
MBD93030 | NI-SSF SUBSTAGE FOR FOCUSING NOSEPIECE |  |  | NISSF:1:1:P | FNC:1:1:H | NI3PSA¦NISA¦NIZDECK:1:1:H
MBD93010 | NI-SSR SUBSTAGE FOR ROT/MOTOR STAGE |  |  | NISSR:1:1:P | NICH¦NICUDE:1:1:H | NISE¦NIECS¦StageH101¦NIUCS:1:1:H
MBD93000 | NI-SS SUBSTAGE |  | Requires Condenser Carrier MBD94000 | NISS:1:1:P | NICH:1:1:H | NISSStage:1:1:H
MBD94000 | NI-CH CONDENSER HOLDER |  | Requires stage carrier | NICH:1:1:P | NiCondenser:1:1:H
MBC91010 | NIE-CSRR2 ROTATING CERAMIC STAGE R W/HDR |  | Discontinued - Limited inventory. Use MBC91016 | NIECS:1:1:P | NIEStageCompatibility:1:1:P
MBC91016 | NIE-CSRR2 ROTATING CERAMIC STAGE R W/HDR |  | Replaces MBC91010 and MBC91015 | Alias:MBC91010
MBC91020 | NIE-CSLR2 ROTATING CERAMIC STAGE L W/HDR |  |  | Alias:MBC91010
MBC96100 | NI-S-E MOTORIZED XY STAGE |  | Requires joystick or ergo and power cord | NISE:1:1:P | NIEStageCompatibility:1:1:P
MBC91110 | C-SR2S RIGHT HANDLE STAGE WITH 2S HOLDER |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91136 | C-CSR2S CERAMIC R HANDLE STAGE W/2S HDR |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91115 | C-CSR1S RGHT HANDLE CERAMIC STG 1S HLDR |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91146 | C-CSR1S RGHT HANDLE CERAMIC STG 1S HLDR |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91120 | C-SR25 LEFT HANDLE STAGE W/2S HOLDER |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91155 | C-SL2S LEFT HANDLE STAGE W/2S HOLDER |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91126 | C-CSL2S LEFT HAND CERAMIC STAGE W/2S HDR |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91166 | C-CSL2S LEFT HAND CERAMIC STAGE W/2S HDR |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91130 | C-CSR RIGHT HANDLE CERAMIC STAGE |  | Requires Slide Holder and MBD93000 Stage Carrier | NISSStage:1:1:P | CH2L:1:1:H | NILStageCompatibility:1:1:P
MBC91176 | C-CSR RIGHT HANDLE CERAMIC STAGE |  | Requires Slide Holder and MBD93000 Stage Carrier | NISSStage:1:1:P | CH2L:1:1:H | NILStageCompatibility:1:1:P
92620 | I SERIES PLAIN STAGE |  | Requires MBD93000 Stage Carrier | NISSStage:1:1:P |  | NILStageCompatibility:1:1:P
MBC91046 | NIU-CSLR2 ROTATING CERAMIC STAGE L W/HDR |  |  | NIUCS:1:1:P |  | NILStageCompatibility:1:1:P
MBC91030 | NIU-CSRR2 ROTATING CERAMIC STAGE R W/HDR |  | Requires MBD93010 Stage Carrier | NIUCS:1:1:P |  | NILStageCompatibility:1:1:P
MBC91036 | NIU-CSRR2 ROTATING CERAMIC STAGE R W/HDR |  | Requires MBD93010 Stage Carrier | NIUCS:1:1:P |  | NILStageCompatibility:1:1:P
MDC83000 | P-GS2 CIRCULAR GRADUATED STAGE |  | Requires MBD65003 LV-SAD Stage Adapter plus MBD93010 Stage Carrier.
MBD65003 | LV-SAD STAGE ADAPTER
MBC95020 | C-H2L SPECIMEN HOLDER 2L |  |  | CH2L:1:1:P
MBC95035 | C-H2L SPECIMEN HOLDER 2L |  |  | CH2L:1:1:P
MBL99005 | NI-CUD UNIVERSAL CONDENSER DRY-NC |  |  | NiCondenser:1:1:P | NICUDModule:1:7:H
MBL99800 | NI-CUD-E MOTORIZED UNIVERSAL COND DRY |  |  | NICUDE:1:1:P | NICUDModule:1:7:H | NIECondenserCompatibility:1:1:H
MBL78600 | D-CUD UNIVERSAL CONDENSER DRY |  | Older version - Limted quantity available
MBB75370 | C-SP SIMPLE POLARIZER | Polarizer |  | NiBasePolarizer:1:1:P
MBB75380 | C-TP POLARIZER W/RED TINT PLATE F/GOUT | Polarizer |  | NiBasePolarizer:1:1:P
MBN74940 | D-DP DIC ROTATABLE POLARIZER | Polarizer |  | NiBasePolarizer:1:1:P
MBN91930 | NI-PT POLARIZER TURRET | Polarizer | Fits on Ni-E Base | NiBasePolarizer:1:1:P | FNPolarizer:1:1:H
MBN71946 | FN-P POLARIZER | Polarizer |  | FNPolarizer:1:1:P
MBN71960 | FN-IRP IR POLARIZER | Polarizer |  | FNPolarizer:1:1:P
MBE95205 | NI-LH PRECENTERED LAMPHOUSE-NC |  |  | LamphouseU:1:1:P | SHADI:0:1:H | NICTLA:1:1:H
MBE75235 | FN-LH PRECENTERED LAMPHOUSE-NC |  |  | LamphouseU:1:1:P | SHADI:0:1:H | NICTLA:1:1:H
MBE65800 | C-LL HIGH COLOR RENDER LED LIGHT SOURCE |  | Recommended for transmitted light microscope applications. If Epi-fluorescence is added to the microscope - a transmitted light shutter is needed due to autofluorescence from the white phospher LEDs | LamphouseU:1:1:P | SHADI:0:1:H | NICTLA2:1:1:H
MBD92010 | NI-AMSS ARM SPACER 49MM |  | Requires contact or standard arm | NIAMS:1:1:P | NIFNA:1:1:H | FNNosepiece:1:3:H | NIAMSCompatibility:1:1:P
MXB22379 | NI-AMSL ARM SPACER 74MM |  | Requires contact or standard arm | Alias:MBD92010
MBC94100 | NI-3PSA ADAPTER FOR FN1 STANDARD STAGE |  | Requires MBD93030 substage | NI3PSA:1:1:P | FN3PS2:1:1:H
MBC74001 | FN-3PS2 RECTANGULAR STAGE |  | Requires MBC94100 stage adapter | FN3PS2:1:1:P
MBC74005 | FN-3PS2 RECTANGULAR STAGE-NC |  | Requires MBC94100 stage adapter | FN3PS2:1:1:P
MBC96110 | NI-SA MOTORIZED XY STAGE ADAPTER |  | Requires MBD93030 substage | NISA:1:1:P | NISE:1:1:H
MBL70105 | FN-C LWD CONDENSER-NC |  |  | FNC:1:1:P | FNPT4X¦NICUDModule:1:2:H
MBH72100 | FN-PT4X 4X AUXILIARY LENS
MBB94511 | NI-TT2-E MTRZD QUADROCULAR TILTING TUBE2 | NiLayer | Manual Laser Safe Quad Head for Ax. Necessary for 25mm FOV. | NITTE:1:1:P | NIRPZ¦NIRPZE:0:1:H | CTAQ¦CTAQ2¦A1UTT:0:1:H | EyePiece¦EyepieceUW:2:2:H
MBB94510 | NI-TT-E MOTORIZED QUADROCULAR TILT TUBE | NiLayer | FOLLOW UP | Alias:MBB94511
MBB94522 | NI-TT2 QUADROCULAR TILTING TUBE 2 | NiLayer | Manual Laser Safe Quad Head for Ax. Necessary for 25mm FOV. | NITT:1:1:P | NIRPZ:0:1:H | CTAQ¦CTAQ2¦A1UTT:0:1:H | EyePiece¦EyepieceUW:2:2:H
MBB94521 | NI-TT QUADROCULAR TILTING TUBE | NiLayer | FOLLOW UP | Alias:MBB94522
MBB96710 | NI-RPZ-E MOTOR DSC ZOOM PORT F/QUAD TUBE |  |  | NIRPZE:1:1:P | CmountCamera:0:1:H
MBB96721 | NI-RPZ DSC ZOOMING PORT F/QUAD TUBE |  |  | NIRPZ:1:1:P | CmountCamera:0:1:H
MBB96900 | NI-RPS DSC CAMERA SUPPORT PILLAR |  | Requires support plate
MBB96910 | NI-RPSAA DSC ADAPTER A (ANDOR) |  | Requires support pillar
MBB96920 | NI-RPSAB DSC ADAPTER B (PHOTOMETRICS) |  | Requires support pillar
MBB94700 | C-TAQ TUBE ADAPTER F/QUADROCULAR TUBE |  | Requires ISO Adapters | CTAQ:1:1:P | CmountAdapter:1:1:H
MHV40100 | C2-NI-TT ADAPTER FOR QUAD TILTING TUBE
MHV55070 | A1-U-TT-FN1/NI MOUNTING BRIDGE |  | Required | A1UTT:1:1:P | AXScanHead:1:1:H
MBB93800 | C-TE2 ERGONOMIC BINOCULAR TUBE | NiLayer |  | UprightTube:1:1:P | CTEPF¦CTEP:0:1:H | EyePiece:2:2:H
MBB96800 | C-TEP2 DSC PORT FOR ERGO BINOC TUBE 0.7X |  |  | CTEP:1:1:P | CmountCamera:1:1:H
MBB96855 | C-TEP3 DSC PORT F/ERGO BINOC TUBE 0.55X |  |  | CTEP:1:1:P | CmountCamera:1:1:H
MBB96810 | C-TEPF2 DSC PORT ERGO BINOC FMOUNT 2.5X |  | For Clinical Ergo Heads [MBB93800] | CTEPF:1:1:P | FmountCamera:1:1:H
MBB92106 | C-TB BINOCULAR TUBE-NC | NiLayer |  | UprightTube:1:1:P |  | EyePiece:2:2:H
MBB93115 | C-TT TRINOCULAR TUBE T-NC | NiLayer |  | UprightTube:1:1:P | YTV:1:1:H | EyePiece¦EyepieceUW:2:2:H
MBB93106 | C-TF TRINOCULAR TUBE F-NC | NiLayer |  | UprightTube:1:1:P | YTV:1:1:H | EyePiece¦EyepieceUW:2:2:H
MBB73550 | Y-TV55 TV TUBE FOR C-0.55X RELAY |  | FOLLOW UP: How does this fit F mount and .55 C mount? | YTV:1:1:P | FmountAdapter:1:1:H
MBB73525 | Y-T TV TUBE/C MOUNT ADAPTER A-NC |  |  | YTV:1:1:P | CmountAdapter:1:1:H
MBB94800 | C-TAQ2 TUBE ADAPTER FOR QUADROCULAR |  | For Ni/FN Quadrocular heads [MBB94510, MBB94521] | CTAQ2:1:1:P | FmountAdapter:1:1:H
MHV40000 | C2-RA RING ADAPTER FOR NI
MBB74700 | QUADROCULAR EYEPIECE TUBE ADAPTER
MBB75710 | C-ER EYELEVEL RISER | NiLayer |  | NiIntermediateTube:1:1:P | UprightTube:1:1:H
MBB74100 | Y-IDP DOUBLE PORT 55/45 | NiLayer | Recommend using Ci type body tubes when using double ports | NiIntermediateTube:1:1:P | UprightTube:1:1:H
MBB74105 | Y-IDP DOUBLE PORT 0/100 | NiLayer | Recommend using Ci type body tubes when using double ports | NiIntermediateTube:1:1:P | UprightTube:1:1:H
MBB75500 | Y-IM MAG. CHANGER
MAK31100 | CFI UW EYEPIECE 10X M |  |  | EyepieceUW:1:1:P
MXA22004 | CFI NOSEPIECE DUST CAP
MEF51001 | TI-100WRC 100W LAMPHOUSE REMOTE CABLE
MBF72230 | NI-FIA TRANSMITTED LIGHT FIBER APAPTER |  |  | LamphouseU:1:1:P | LLG:1:1:H
MXK37170 | IQV 100W HALOGEN LAMP FOR LS-DWL-N
MBN66730 | 25MM GIF FILTER
MBE65700 | LV-LL LED PRECENTERED LAMPHOUSE |  | Recommended for transmitted light microscope applications. If Epi-fluorescence is added to the microscope - a transmitted light shutter is needed due to autofluorescence from the white phospher LEDs | LamphouseU:1:1:P |  | NICTLA:1:1:H
MBH99000 | NI-CALN1 2-4X AUXILIARY LENS F/POS N1 |  | Required for 2x and 4x objectives | NICUDModule:1:1:P
MBH70200 | D-C 2-4X AUXILIARY LENS |  | Required for 2x and 4x objectives | NICUDModule:1:1:P
MBH70210 | D-C PH MODULE PH-1 DRY |  |  | NICUDModule:1:1:P |  | PH1Module:0:6:P
MBH70220 | D-C PH MODULE PH-2 DRY |  |  | NICUDModule:1:1:P |  | PH2Module:0:6:P
MBH70230 | D-C PH MODULE PH-3 DRY |  |  | NICUDModule:1:1:P |  | PH3Module:0:6:P
MBH72300 | D-C DIC MODULE N1 DRY |  |  | NICUDModule:1:1:P |  | DICPrismN1:0:6:P
MBH72310 | D-C DIC MODULE N2 DRY |  |  | NICUDModule:1:1:P |  | DICPrismN2:0:6:P
MBH72330 | D-C DIC MODULE NR DRY |  |  | NICUDModule:1:1:P |  | DICPrismNR:0:6:P
MBH70250 | D-C DARK FIELD RING |  | Fits in N2 or Aux Position | NICUDModule:1:1:P
MBL78700 | D-CUO DIC CONDENSER OIL |  |  | NICUDModule:1:1:P | DICNROil:0:1:H | DICPrismNR:0:6:P
MBH72340 | D-C DIC MODULE NR OIL |  |  | DICNROil:1:1:P
MBL11300 | FLIP OUT CONDENSER |  | Fixed Substage Only! | NiCondenser:1:1:P
MBL71305 | CC SWING-OUT ACHROMAT CONDENSER 1-100XNC |  | White Top | NiCondenser:1:1:P
MBL71505 | C-C SLIDE ACHRO CONDENSER 2-100X-NC |  |  | NiCondenser:1:1:P
MBL12000 | DARKFIELD CONDENSER OIL F/XY |  |  | NiCondenser:1:1:P
MBL12010 | DARKFIELD CONDENSER DRY F/XY |  |  | NiCondenser:1:1:P | CCEL:0:1:H
MXA20740 | C-CEL COND BOOSTER EXP LENS |  |  | CCEL:1:1:P
MBL71405 | CC ACHROMAT/APLANAT CONDENSER NA 1.4-NC |  |  | NiCondenser:1:1:P
MBL73105 | CC PHASE TURRET CONDENSER NA 0.9-NC |  | Fixed Substage Only!
MBL16105 | LWD ACHROMAT CONDENSER-NC |  |  | NiCondenser:1:1:P
MBB75350 | C-ISA INTER TUBE W/ SIMPLE ANALYZER | NiLayer | Intermediate tube (25mm High) | NiIntermediateTube:1:1:P | UprightTube:1:1:H
MBN71950 | FN-IR/FA FL/DIC IR ANALYZER | Analyzer | For Use with MBE74100 D-FL UNIVERSAL EPI-FLUOR ILLUMINATOR Only | DFL2Slider:1:1:P
MBN71955 | FN-IR/ISA DIC IR ANALYZER | Analyzer | Intermediate tube (25mm High) | NiIntermediateTube:1:1:P | UprightTube:1:1:H
MBN71941 | D-FA/IR FL/DIC IR ANALYZER | Analyzer | For Use with MBE74100 D-FL UNIVERSAL EPI-FLUOR ILLUMINATOR Only | DFL2Slider:1:1:P
MBE94500 | NI-FLT6 EPI-FLUORESCENCE CUBE TURRET | NiLayer |  | NIFLT6:1:1:P | UprightTube¦NITT:1:1:H | NIFLEI¦NIBPU:1:1:H | NIFAFL¦NIAST:0:1:H | FilterCube:1:6:H
MBE94400 | NI-FLT6-I INTELLI EPI-FLUOR CUBE TURRET | NiLayer |  | NIFLT6I:1:1:P | UprightTube¦NITT¦NITTE¦NIFLT6I¦NIFLT6E:1:1:H | NIFLEI¦NIBPU:1:1:H | NIFAFL¦NIAST:0:1:H | FilterCube:1:6:H
MBE94300 | NI-FLT6-E MOTOR EPI-FLUOR CUBE TURRET | NiLayer |  | NIFLT6E:1:1:P | UprightTube¦NITT¦NITTE¦NIFLT6I¦NIFLT6E:1:1:H | NIFLEI¦NIBPU:1:1:H | NIFAFL¦NIAST:0:1:H | FilterCube:1:6:H
MBE94100 | NI-FLEI EPI-FLUORESCENCE ATTACHMENT |  | Discontinued. Limited stock. Check availability. Replace with MBE94150 plus MXA22205, or MBE94150 plus MXA22198, or MBE94150 plus MBF83000 (DS21-014). NOTE: MBE94100 does not contain a POL slot | NIFLEI:1:1:P | DLED¦FIBA:1:1:H
MBE94150 | NI-FLEI-2 EPI-FLUORESCENCE ATTACHMENT |  | Direct Connect to Nikon D-LEDI LED Illuminator MBF83000. Use MXA22205 C‒FIBA Adapter For Fiber Light with 5mm Core Diameter LLG. Use MXA22198 D-LEDI Retrofit Adapter for HMX Connection. | NIFLEI:1:1:P | DLED¦FIBA:1:1:H
MBE94200 | NI-BPU BACK PORT UNIT |  | FOLLOW UP | NIBPU:1:1:P
MBE74100 | D-FL UNIVERSAL EPI-FLUOR ILLUMINATOR | NiLayer | Discontinued. Stock limited to consignment inventory. Replace with MBE74150 plus MXA22205, or MBE74150 plus MXA22198 or MBE74150 plus MBF83000 (DS21-014) | Alias:MBE74150
MBE74150 | D-FL-2 U-EPI FLUORESCENSE ATTACHMENT | NiLayer | Direct Connect to Nikon D-LEDI LED Illuminator MBF83000. Use MXA22205 C‒FIBA Adapter For Fiber Light with 5mm Core Diameter LLG. Use MXA22198 D-LEDI Retrofit Adapter for HMX Connection. | NIFLT6:1:1:P | UprightTube¦NITT:1:1:H | DLED¦FIBA:1:1:H | NIFAFL¦NIAST:0:1:H | FilterCube:1:6:H | DFL2Slider:0:1:H
MBE49100 | D-FLE FTR BLOCK F/BF-DIC-SMPLE POL ILLUM |  | Not for use with Ni-PAU arm
MBE49200 | D-FLD FTR BLOCK FOR EPI DARK FIELD ILLUM |  | Not for use with Ni-PAU arm
MBN71956 | A1-FN1-IR ANALYZER BLOCK IR
MBE64100 | LV-PAB PA BLOCK |  | Not for use with Ni-PAU arm
MBN71930 | E-FB EXCITATION BALANCER
MBN66921 | YM-PO POLARIZER FOR U-EPI/L-UEPI
MBE75060 | C-FX EPI-FL CENTERING TOOL
MBE94800 | NI-PAU NI PHOTOACTIVATION UNIT
MXA22118 | NI-DRS DSC ROTATION STOPPER
91158 | 50/50 MIRROR FOR PA-GFP ILLUM
91159 | FULL ENHANCED (AG) MIRROR TIRF & PA-GFP
91160 | FULL MIRROR (AL) TIRF & PA-GFP
MEE58015 | TI-PAU DICHROIC MIRROR PA-GFP
MEE59115 | TI-TIRF-HM 60/40 MIRROR
MXA22089 | TI SHUTTER TRIGGER CABLE/HGFIE
MBE25305 | LH-M100C-1 LAMPHOUSE HMX-3B-NC
MBE25315 | LH-M100CB-1 LAMPHOUSE HMX-4-NC
MEE45900 | TE DOUBLE LAMPHOUSE ADAPTER
78589 | MERCURY BULB, USH102DH |  | For Older HBO Power Supplies with Black Front Panel Only
78590 | HBO100W/12 LAMP F/MBF12321 HBO103W/2
MBF33345 | MERCURY SOCKET 100W BLACK-NC
MBF12321 | HBO 100W POWER SUPPLY 100-240V
MBE75015 | CFC EPI-FL COLLECTOR LENS-NC
MBE15003 | COLLECTOR LENS F/QUARTZ
MBC95055 | NI-SH-D PETRI DISH HOLDER
MXA22088 | TI SHUTTER TRIGGER CABLE/SUTTER
MBE96900 | NI-SHAEP-U MTR SHUTTER ADPTR F/UP EPI-FL
MBE96920 | NI-SHCS MOTORIZED SHUTTER CABLE SHORT
84125 | HALOGEN LAMP 12V-100W L.L.
175961 | 27MM RETICLE 10MM SCALE
175966 | 27MM 90 DEG CROSSLINE WITH 10MM SCALE
175962 | 27MM RETICLE 1MM SQUARE
MBM11100 | STGE MCRMTR A 1MM IN 100 DIV
MXA22048 | NYLON DUST COVER H550L (FOR 80I/90I)
MXA22049 | NYLON DUST COVER H600L (50I/55I/EPI-FL)
92623 | NIE FIXING METAL PLATES
84372 | SFTU-1 OPTIONAL MTG KIT FOR CI-NI |  | To
MXA22122 | CI-IRC IR CUT FILTER FOR CI
MBN91921 | C-AS ANALYZER FOR GOUT |  | Installs in Nosepiece Slot
MBB75360 | C-IA INTERMEDIATE TUBE W/ ANALYZER-GOUT |  | Intermediate Tube (25mm High)
MBN74970 | D-AA ANALYZER FOR ASBESTOS |  | Install in MBB75770 Iintermediate Tube for Asbestos
MBB75770 | D-IMA INTERMEDIATE TUBE FOR ASBESTOS |  | Intermediate Tube (25mm High) Accepts MBN74970 Analyzer for Asbestos ONLY
MBB75400 | Y-IDT DRAWING TUBE

## Ci
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
MBA940SP | CI-S BASIC ERGO BINOCULAR PACKAGE |  | Consists of Ci-S Main Body with 6V/30W Halogen Illumination, C-TE2 Ergonomic Tilting and Telescopic Binocular Head, 10x/FN22 Eyepieces (x2), 10mm High Eyecups (x2), C-N Sextuple Nosepiece ESD, C-CSR1S Right Handle Ceramic Stage with 1S Holder, Dust Cover H550S. NOTE: Condenser, Objectives and Power Cord Sold Separately.
MBA950QA | CI-L PLUS BASIC ERGO BINOC PCK-1S HLDR |  | Consists of the Ci-L plus Main Body, Intelligent 6-Place Nosepiece with Analyzer Slot, AC Adapter, 10x/FN 22 Eyepieces (x2), Eyeguards (x2), Ergonomic Binocular Eyepiece Tube, Right Hand Ceramic Stage with a 1 Slide Sample Holder, Nosepiece Caps (x3), Nylon Dust Cover and Instruction Manual. NOTE: Condenser, Objectives and Power Cord Sold Separately. Optional USB to Mini-B cable req'd for programming objectives into LIM system.
MBA940BB | CI-S BASIC BINOCULAR HEAD PACKAGE |  | Consists of Ci-S Main Body with 6V/30W Halogen Illumination, C-TB Binocular Head with Siedentorf Eyepiece Tube, 10x/FN 22 Eyepieces (x2), C-N Sextuple Nosepiece ESD, C-SR2S Right Handle Stage with Double Slide Holder (2S), Dust Cover H550S. NOTE: Condenser, Objectives and Power Cord Sold Separately.
MBA950RA | CI-L PLUS BASIC BINOC PCK-2S HLDR |  | Consists of the Ci-L plus Main Body, Intelligent 6-Place Nosepiece with Analyzer Slot, AC Adapter, 10x/FN 22 Eyepieces (x2), Eyeguards (x2), Binocular Eyepiece Tube, Right Hand Ceramic Stage with a 2 Slide Sample Holder, Nosepiece Caps (x3), Nylon Dust Cover and Instruction Manual. NOTE: Condenser, Objectives and Power Cord Sold Separately. Optional USB to Mini-B cable req'd for programming objectives into LIM system.
MBA950SA | CI-L PLUS BASIC ERGO BINOC PCK-1S W/FI3 |  | Consists of the Ci-L plus Main Body, Intelligent 6-Place Nosepiece with Analyzer Slot, AC Adapter, 10x/FN 22 Eyepieces (x2), Eyeguards (x2), Ergonomic Binocular Eyepiece Tube, Right Hand Ceramic Stage with a 1 Slide Sample Holder, 0.55x DSC Camera Port, DS-Fi3 Camera, AC Adapter for Camera, Nosepiece Caps (x3), Nylon Dust Cover and Instruction Manual. NOTE: Condenser, Objectives, Power Cords (x2) and USB 3 Cable for Camera Sold Separately. Optional USB to Mini-B cable req'd for programming objectives into LIM system.
77012831 | 6FT USB 2.0 A TO MINI-B 5-PIN CABLE |  | USB2.0 to 5-pin USB-mini (male-to-male) cable required when connecting the Eclipse Ci-L Plus to a computer for objective registration into the LIM system via Ni Tools or NIS-Elements.
MBA94010 | ECLIPSE CI-S MAIN BODY-NC |  | Eclipse Ci-S Main Body with 6V/30W Halogen Lamp, Re-Focusing Mechanism, Built-In ND4 and ND8 Filters, Fixed Diffuser, Field Diaphragm, Capture Button, Adjusting Tool (1 Piece). NOTE: Power Cord Sold Separately.
MBA95010 | ECLIPSE CI-L PLUS MAIN BODY |  | Eclipse Ci-L plus Main Body with white LED illumination, Image Capture Button, LIM (Light Intensity Management) System, Intelligent 6-Place Nosepiece with Analyzer Slot, Nosepiece Caps (x3), AC Adapter, and Instruction Manual. NOTE: Power Cord Sold Separately. Optional USB to Mini-B cable req'd for programming objectives into LIM system.
MBA93000 | ECLIPSE CI-E MAIN BODY-NC |  | Eclipse Ci-E Main Body with LED Illumination and Motorized Sextuple Nosepiece with Slot and Controller, Focusing Mechanism, Fixed Diffuser, Field Diaphragm, Capture Button, Adjusting Tool (1 Piece). NOTE: Power Cord Sold Separately.
MXA22129 | C-FK EXTND FINE FOCUS KNOB FOR CI/NI-NC
MXA22113 | C-CTC CAMERA TRIGGER CABLE U3/L3
MAK10150 | CFI 15X EYEPIECE F.N. 14.5 MM
MBP91705 | CI-NS SPACER FOR NOSEPIECE-NC |  | Compatible with "all" Ci stands. Must Quote 999-161 for installtion on Ci-E and CIL-Plus | CiSpacer:1:1:P
MBL99300 | CI-C-E MOTORIZED SWING-OUT CONDENSER
MBL71105 | CC ABBE CONDENSER N.A. 0.9 NC
MBL71300 | CC SWING-OUT ACHROMAT CONDENSER 1-100X |  | Older version - limited demo stock
MBL11305 | SWING-OUT ACHRO CONDENSER 2-100X CH |  | Formally referred to as "Flip-Out" type
MBL71205 | CC-AR ACHROMAT CONDENSER N.A.0.80-NC
MCN77100 | E2-F-BF BLUE LIGHT BLOCKING/CUT FILTER
MXK37554 | MC-WIO COLOR BALANCE FILTER 45MM
MBE94350 | CI-FL-2 EPI-FLUORESCENCE ATTACHMENT |  | Direct Connect to Nikon D-LEDI LED Illuminator MBF83000. Use MXA22205 C‒FIBA Adapter For Fiber Light with 5mm Core Diameter LLG. Use MXA22198 D-LEDI Retrofit Adapter for HMX Connection.
MBB75100 | Y-THFF TH ATTACHMNT FF
MBB75110 | Y-THSSB TH ATTACHMNT SSB
MXA22065 | COVER TYPE 117
MBB75120 | Y-THM MAIN TEACHING UNIT |  | Do not pair with a Ci Main Body in a Teaching Configuration greater than 5-Heads. Can be paired with an Ni Main Body for Teaching Configurations up to 18-Heads.
MBB75131 | Y-THR-L TEACHING UNIT SSA TYPE-L
MXA22064 | COVER TYPE 116-MULTI VIEWING
MBB75150 | Y-THS SUPPORT SS UNIT

## Ji
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
97500 | ECLIPSE JI ASSAY MICROSCOPE PACKAGE | Base | Eclipse Ji Assay Microscope Package Consisting of: Eclipse Ji Digital Inverted Microscope, Ji-CN Controller, Ji-SC Side Port Cover, D-LEDI2 Fluorescence LED Illumination System, AC Adapter for D-LEDI, C-FLL LFOV Quad D/F/T/Cy5 LED Filter Set, NIS-Elements AR, NIS-Elements Module: Jobs Acquisition, NIS-Elements Module: General Analysis 2 and 3, NIS-Elements: 6D Imaging, NIS-SE Smart Experiment Assays, CFi60 Plan Apochromat Lambda D 4x, CFi60 Plan Apochromat Lambda D 10x, CFi60 Plan Apochromat Lambda D 20x, 2 each USB Cable, 2 each Power Cord. | AddPart:MFA92000,MFF92000,MFD92000,MBF83200,MQF52057,96378,MQS31000,MQS43130,77010604,MQS42560,MQS44000,MRD70040,MRD70170,MRD70270,97050,97050,79035,79035
97505 | ECLIPSE JI RESEARCH MICROSCOPE PACKAGE | Base | Eclipse Ji Research Microscope Package Consisting of: Eclipse Ji Digital Inverted Microscope, Ji-CN Controller, NIS-SE Smart Experiment Assays, USB Cable, Power Cord. | AddPart:MFA92000,MFF92000,MQS44000,97050,79035
MFC92000 | C-HS HOLDER ADAPTER SET (35MM & 60MM) |  |  | StageUInsert:0:1:P
MFD92000 | JI-SC SIDE PORT COVER |  |  | JiSidePortCover:1:1:P
MFD92010 | JI-SC-AX SIDE PORT COVER FOR AX |  |  | JiSidePortCover:1:1:P
MFD92020 | JI-FP FIXING PLATE | FixingPlate | Required for use with Ji Confocal Application on VIT | FixPlateJi:1:1:P
77025345 | JIX STG TP INCUBATION SYS F/100% CO2 |  | System includes: Incubator Chamber, Temperature Control Unit with Gas Function, Dish Holder set for Wellplate, Chamber Slide and Chambered Coverglass, Sample Feedback Regulation, Tokai Hit Software (Data Logging, PC Regulation) and Stage Adapter for Ji XY Motorized Stage. NOT COMPATIBLE WITH QUEENSGATE PIEZO. | JiIncubator:1:1:P | TokaiExtension:1:1:H
77025346 | JIX STG TP INCUBATION SYS F/5% CO2 |  | System includes: Incubator Chamber, Temperature Control Unit with Gas Function, Dish Holder set for Wellplate, Chamber Slide and Chambered Coverglass, Sample Feedback Regulation, Tokai Hit Software (Data Logging, PC Regulation) and Stage Adapter for Ji XY Motorized Stage. NOT COMPATIBLE WITH QUEENSGATE PIEZO. | JiIncubator:1:1:P
77025382 | 1.0M EXTENSION CABLE F/TOKAI STX INCUB |  | Required for elongating Tokai STX style chamber cable for JI use, if not using JIX specific chambers (77025345 or 77025346) | TokaiExtension:1:1:H
97522 | Ji VIBRATION ISOLATION FEET SET |  | Recommended for use with Ji on lab bench. Not for use with Ji Confocal | JiFeet:1:1:P
MFA92000 | ECLIPSE JI DIGITAL INVERTED MICROSCOPE | LaserSafety,Ji |  | SidePortAdapter:0:1:H | FixPlateJi¦JiFeet:0:1:H | InterlockRound:0:1:H | AddPart:JBAM-Ti2SidePort | StageEInsert:1:6:H | JiIncubator:0:1:H | JiController:1:1:H | JiSidePortCover:0:1:H | CFI60Objective:1:6:H | FilterCube¦FilterCubeLFOV:0:6:H | FIBA¦DLED:1:1:H
MFF92000 | JI-CN CONTROLLER |  |  | JiController:1:1:P | USBCord:1:1:H | PowerCord:1:1:H

## SMZ18¦25
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
Stands
MNA54000 | SMZ-25 ZOOM STEREO MICROSCOPE | Base | Requires Control Box A | SMZ25ZoomingBody:1:1:P | SMZRelayBox¦SMZFocusRelay:1:1:H | SMZBoxA:1:1:H | SMZObjective:1:4:H | SMZEpiAttachment:0:1:H | SMZTrinoc:1:1:H
MNA53000 | SMZ-18 ZOOM STEREO MICROSCOPE | Base |  | SMZ18ZoomingBody:1:1:P | SMZRelayBox¦SMZFocusRelay¦SMZRelayCable:1:1:H |  | SMZObjective:1:4:H | SMZEpiAttachment:0:1:H | SMZTrinoc:1:1:H
Control Box
MNF55030 | SMZ-25 P2-CTLA CONTROL BOX A |  |  | SMZBoxA:1:6:P | CACAdapter:1:1:H | SMZRC:0:1:H
MNF55010 | SMZ-25 P2-RC REMOTE CONTROLLER |  |  | SMZRC:1:1:P
MNF55040 | SMZ-18 P2-CTLB CONTROL BOX B |  |  | SMZBoxB:1:6:P | CACAdapter:1:1:H
Body Tubes
MNB53510 | SMZ-25/18 P2-TERG 100 TRINOC TILT TUBE-N |  | Requires Tube adapter for ISO | SMZTrinoc:1:1:P | SMZEyepiece:2:2:H | FmountAdapter:0:1:H
MNB53520 | SMZ-25/18 P2-TERG 50 TRINOC TILT TUBE-NC |  | Requires Tube adapter for ISO | Alias:MNB53510
MNB53560 | SMZ-25/18 P2-TL100 TRINOC TUBE 100-NC |  | Requires Tube adapter for ISO | Alias:MNB53510
Eyepieces
MMK30102 | C10XB WIDEFIELD EYEPIECE - NC |  | 0 | SMZEyepiece:1:1:P | SMZEyegaurd:0:1:H | SMZReticle:0:1:H
MMK30155 | C 15X WIDE EYEPIECE-CH |  | 0 | SMZEyepiece:1:1:P | SMZEyegaurd:0:1:H
MMK30205 | C 20X WIDE EYEPIECE-CH |  | 1.3x Reticle Mag Factor | SMZEyepiece:1:1:P | SMZEyegaurd:0:1:H
MMK30300 | C 30X EYEPIECE |  | 1.4x Reticle Mag Factor | SMZEyepiece:1:1:P
MXA25008 | HORN-TYPE GUARD F/CW-10/15/20/30X-EYEPCS |  | Shaped to sheild eye from room light | SMZEyegaurd:0:1:H
MXA25014 | RUBBER RING FOR C-W10X AIR TIGHT - NC |  | Replacement
Eyepiece Reticles
MMM10100 | C-FC RETICLE C-CROSSLINE-25MM |  | Crossline w/o any division markings in a reticle holder, 25mm diameter. Note: For MMK30101, MMK30102 eyepieces only | SMZReticle:1:1:P
MMM10110 | C-FM RETICLE M-CROSSED MICROMETER-25MM |  | Crossline with division markings (100div/10mm) on both the x and y axis in a reticle holder, 25mm diameter. NOTE: For MMK30101, MMK30102 eyepieces ony. | SMZReticle:1:1:P
83074 | 25MM RETICLE - 10MM SCALE - 400 SQUARES |  | Discontinued. Limited Inventory. | SMZReticle:1:1:P
83075 | 25MM RETICLE - CROSSLINE - 100 DIVISIONS |  |  | SMZReticle:1:1:P
Objective Carriers
MND54020 | SMZ-25/18 P2-FM FOCUS MOUNT ADAPTR-NC |  |  | SMZFMA:1:1:P | SMZ18ZoomingBody¦SMZ25ZoomingBody:1:1:H
MNP51500 | SMZ-25/18 P2-RNI2 INTELLIGENT NOSEPC-NC |  |  | SMZFMA:1:1:P | SMZ18ZoomingBody¦SMZ25ZoomingBody:1:1:H | SMZRelayBox¦SMZFocusRelay:1:1:H
Objectives
MNH55050 | SMZ-25/18 P2-SHR PLAN APO 0.5X OBJ |  | Ring Light Diameter 81mm | SMZObjective:0:1:P | SMZQWP0.5X:0:1:H | SMZOBJAttachment:0:1:H
MNH55100 | SMZ-25/18 P2-SHR PLAN APO 1X OBJ |  | Ring Light Diameter 81mm | SMZObjective:0:1:P | SMZQWP1X:0:1:H | SMZOBJAttachment:0:1:H
MNH55160 | SMZ-25/18 P2-SHR PLAN APO 1.6X OBJ |  | Not for use with Epi Coax or Ring LED | SMZObjective:0:1:P
MNH55200 | SMZ-25/18 P2-SHR PLAN APO 2X OBJ |  | Not for use with Epi Coax or Ring LED | SMZObjective:0:1:P
Focusing Columns for P2 Bases
MND52020 | SMZ-25/18 P2-FU MANUAL FOCUS UNIT-NC |  | Requires P2 base and MNF55050 Relay Box for intelligence or motorization of other components | SMZFocusUnit:1:1:P | SMZFMA:1:1:H
MND52010 | SMZ-25/18 P2-MFU MOTORIZED FOCUS UNIT-NC |  | Requires P2 base and Control Box A | SMZFocusUnit:1:1:P | SMZFMA:1:1:H | SMZBoxA:1:1:H | AddPart:JBAM-SMZFocusRelay | SMZFocusRelay:1:1:H
MNF55050 | SMZ-18 P2-RLY RELAY BOX |  |  | SMZRelayBox:1:3:P | SMZBoxA¦SMZBoxB:1:1:H
Bases
MND51010 | SMZ-25/18 P2-PB PLAIN BASE-NC |  | Requires MND52010 or MND52020 P2 Focus Column | SMZFocusUnit:1:1:H | SMZFiberHolder:0:1:H | SMZTiltingStage¦SMZSlidingStage¦SMZPSXYStage:0:1:H | SMZH101Adapter:0:1:H | ThermoPlateP2:0:1:H
MND51510 | SMZ-25/18 P2 DBL LED DIA ILLUM BASE-NC |  | Requires MND52010 or MND52020 P2 Focus Column | Alias:MND51010 | SMZPolAttachment:0:1:H | ACAdapter¦SMZBoxA¦SMZBoxB:1:1:H |  | ThermoPlateDia:0:1:H
MND51610 | SMZ-25/18 P2-DBF FIBER DIA ILLUM BASE-NC |  | Requires either MND52010 or MND52020 P2 Focus Column & LED Source or Halogen (100W or less) | Alias:MND51010 | SMZPolAttachment:0:1:H | CFLED:1:1:P |  | ThermoPlateDia:0:1:H
MND51030 | SMZ18/1270/800N P-PS32 PLAIN STAND-NC |  | Requires MND54010 Focus Mount. | SMZFMDN¦SMZFMCN:1:1:H | SMZFiberHolder:0:1:H | SMZTiltingStage¦SMZSlidingStage¦SMZPSXYStage:0:1:H | SMZH101Adapter:0:1:H | SMZPillar:0:1:H
MND51035 | SMZ18/1270/800N P-PS32 PLAIN STAND 2-NC |  | Stand with 32mm column and 180mm opal/black stage plate. Requires MND54000 (C-FMCN) focus mount. | Alias:MND51030
MND51520 | SMZ18/1270/800N P-DSL32 LED DIA STND-NC |  | Requires MND54010 Focus Mount. AC adapter (MQF52056) and power cord (79035) need to be ordered separately. | Alias:MND51030 | SMZPolAttachment:0:1:H | ACAdapter¦SMZBoxA¦SMZBoxB:1:1:H |  | ThermoPlateDia:0:1:H
MND51620 | SMZ18/1270/800N P-DSF32 FIBR DIA STND-NC |  | Requires MND54010 Focus Mount. | Alias:MND51030 | SMZPolAttachment:0:1:H | CFLED:1:1:P |  | ThermoPlateDia:0:1:H
Base components
MND54010 | SMZ-18 P2FMDN FOCUS MOUNT-NC |  |  | SMZFMDN:1:1:P | SMZ18ZoomingBody:1:1:H
MND43025 | C-EP EXTENSION PILLAR CH |  |  | SMZPillar:0:1:P
MNF55060 | SMZ-18 P2-RLYC RELAY CONTROL CABLE |  |  | SMZRelayCable:1:1:P | SMZBoxB:1:1:H
Epi Fluorescence Illuminator
MNE54312 | SMZ-25/18 P2-EFLM2 MOTOR EPI FLUR ATTACH |  | Requires Control Box A, No Internal Shutter. NOTE: Accepts direct attachment of the D-LEDI Fluorescence LED Illumination System and also comes with a fiber adpater for applications requiring a Fiber Illuminator. The D-LEDI LED Fluorescence Illumination System is sold separately. | SMZEpiAttachment:1:1:P | DLED¦LLG:1:1:H | SMZEpiCube:1:4:H | SMZRelayBox¦SMZFocusRelay:1:1:H | SMZBoxA:1:1:H
MNE54310 | SMZ-25/18 P2-EFLM MOTOR EPI FLUOR ATTACH |  |  | Alias:MNE54312
MNE54320 | SMZ-25/18 P2-EFLI2 INTEL EPI FLUR ATTACH |  |  | Alias:MNE54322
MNE54322 | SMZ-25/18 P2-EFLI2 INTEL EPI FLUR ATTACH |  | Requires Control box A or B for cube position, No Internal Shutter. NOTE: Accepts direct attachment of the D-LEDI Fluorescence LED Illumination System and also comes with a fiber adpater for applications requiring a Fiber Illuminator. The D-LEDI LED Fluorescence Illumination System is sold separately. | SMZEpiAttachment:1:1:P | DLED¦LLG:1:1:H | SMZEpiCube:1:4:H | SMZRelayBox¦SMZFocusRelay:1:1:H
83407 | SMZ-25/18 P2-EFLC DAPI HC FILTER SET |  |  | SMZEpiCube:0:1:P
83408 | SMZ-25/18 P2-EFLC ECFP HC FILTER SET |  |  | SMZEpiCube:0:1:P
83409 | SMZ-25/18 P2-EFLC EGFP LP HC FILTER SET |  |  | SMZEpiCube:0:1:P
83410 | SMZ-25/18 P2-EFLC EGFP BP HC FILTER SET |  |  | SMZEpiCube:0:1:P
83411 | SMZ-25/18 P2-EFLC EYFP HC FILTER SET |  |  | SMZEpiCube:0:1:P
83412 | SMZ-25/18 P2-EFLC DS RED HC FILTER SET |  |  | SMZEpiCube:0:1:P
77075053 | CUSTOM 570 FILTER SET MTED SMZ18 CUBE |  | ET570/20x | SMZEpiCube:0:1:P
83413 | SMZ-25/18 P2-EFLC CY5 HC FILTER SET |  |  | SMZEpiCube:0:1:P
83414 | SMZ-25/18 P2-EFLC DAPI SOLA HC FILTR SET |  |  | SMZEpiCube:0:1:P
77074348 | 49008 MCHRY/TXRD FLTR ST MTD SMZ18/25 CB |  |  | SMZEpiCube:0:1:P
77074696 | 49008 MCHRY/TXRD ST MTD SMZ25/18 CUBE |  |  | SMZEpiCube:0:1:P
MNE54991 | SMZ-25/18 P2-EFLBF FILTER CUBE-BRT FLD |  |  | SMZEpiCube:0:1:P | BRTFLDOC:1:1:P
EPI Coaxial Illuminator
MNB55811 | SMZ-25/18 P2-CI COAXIAL EPI ILLUMINATOR |  |  | SMZEpiAttachment:1:1:P | CFLED:1:1:P
MNB55813 | SMZ-25/18 P2-CI .5X QUARTER LAMBDA PLATE |  |  | SMZQWP0.5X:1:1:P | BRTFLDOC:1:1:H
MNB55812 | SMZ-25/18 P2-CI 1X QUARTER LAMBDA PLATE |  |  | SMZQWP1X:1:1:P | BRTFLDOC:1:1:H
Epi LED Ring Light Illuminator
MNE42710 | SMZ-25/18 P2-FIRL LED RING ILLUM UNIT-2 |  | Requires MQF52056 AC adapter. | SMZOBJAttachment:1:1:P | ACAdapter:1:1:H
Fiber Optic Illuminator
MME49620 | SMZ-25/18 C-FIDH FIBER HOLDER |  | Use with MND51010, MND51510, MND51610 & MND51030, MND51520, MND51620 | FiberCarrier:1:1:P
MME49610 | SMZ-25/18 C-FDF FLEX DBL ARM FIBER ILLUM |  | Requires MME49620 Carrier | CFLED:1:1:P | FiberCarrier:1:1:H
MME49600 | SMZ1270/800N DBL ARM FIBER C-FID2 &PLATE |  |  | CFLED:1:1:P
MNE52600 | SMZ-25/18 P2-FIR FIBER LED RING ILLUM |  |  | CFLED:1:1:P | SMZOBJAttachment:1:1:P
MNF52760 | SMZ-25/18 C-FLED FIBR LED LIGHT SOURCE |  | Requires MQF52056 AC Adapter. Can be used with plastic fiber light guides | CFLED:1:1:H | ACAdapter:1:1:H
NI-150 Light Guides & Accessories
83365 | C-NI-150 FBR OPTIC SORCE W/HEAT FLT 115V |  | Do not use with plastic fiber light guides | DualGooseLight:1:1:H
83364 | C-NI-150 FS LGHT SRCE W/FTSW-ON/OFF CONT |  | Do not use with plastic fiber light guides | DualGooseLight:1:1:H
83359 | NII-LED FIBER OPTIC ILLUMINATOR |  | Do not use with plastic fiber light guides. No power cord. Order separately (79035). | DualGooseLight:1:1:H | PowerCord:1:1:H
83378 | C-FILTER HOLDER
83377 | C-DAYLIGHT CORRECTION FILTER
83376 | C-HEAT FILTER
83371 | C-DUAL GOOSENECK LIGHT PIPES |  | Compatible with 83364, 83365 and 83359 Illuminators. | DualGooseLight:1:1:P | GooseFocusLens:0:2:H
83372 | C-DUAL GOOSENECK LIGHT PIPES |  | Compatible with 83364, 83365 and 83359 Illuminators. | DualGooseLight:1:1:P | GooseFocusLens:0:2:H
83363 | C-FOCUS LENS-GOOSENK LIGHT PIPES-SPRG LD |  | Focusing Lens for 83371 and 83372. | GooseFocusLens:0:1:P | GooseLensCap:0:1:H
83388 | POLARIZING CAPS FOR LIGHT PIPES |  | Polarizing Cap for 83363 Focus Lens. | GooseLensCap:0:1:P
Stages
MMC33000 | TILTING STAGE FOR STEREO MICROSCOPES |  | C-TRS Tilting Stage 0-30º, 122mm Top Surface Area. Accepts a 90mm Opal/Black Stage Plate. The Tilt Base Mounts to Stereo Bases with either 90mm or 180mm Openings. When Mounted to 180mm Stereo Stands, a Separate Opal/Black Stage Plate (MXA20002) is required. | SMZTiltingStage:0:1:P | 90mmStagePlate:0:1:H
MNC53000 | C-SSL DIA SLIDING STAGE |  |  | SMZSlidingStage:0:1:P | SMZDarkFieldUnit:0:1:H | 177mmStagePlate:0:1:H
MND51810 | SMZ-25/18 P-DF2 LED Dark Field Unit 2 |  | P-DF2 Attachable LED Darkfield Base with Intensity Control. AC adapter (MQF52057) and power cord (79035) need to be ordered separately. | SMZDarkFieldUnit:0:1:P | ACAdapter:1:1:H | SMZLightShield:0:1:H
MXA25017 | SMZ-25/18 P-DF LIGHT SHIELDING BOX |  | Blocks stray light from MND51810 (P-DF2 LED DARK FIELD UNIT 2). Sits on table around stand. | SMZLightShield:0:1:P
MNN50920 | SMZ-25/18 P2-POL SIMPLE POL ATTACHMENT |  |  | SMZOBJAttachment:1:1:P
MNC85001 | AZ-SRP ROTATABLE PLATE
Mechanical XY Stage and Accessories
MNC52000 | SMZ-25/18 P-SXY64, 6X4 STAGE-NC |  |  | SMZPSXYStage:0:1:P | P-SXYInsert:1:4:H
MNC85020 | AZ-SDA DISH HOLDER ADAPTER |  |  | P-SXYInsert:1:1:P
MFC35011 | C-HT TERASAKI PLT W65MM PETRI DSH HLD-NC |  |  | P-SXYInsert:1:1:P
MNC85002 | AZ-STA THERMO PLATE ADAPTER |  |  | P-SXYInsert:1:1:P | SMZThermoPlate:1:1:H
84457 | TPI-SX THERMO PLATE F/UPRIGHT MICROSCOPE |  | Tokai Hit | P-SXYInsert:1:1:P
Thermal Stages
84454 | TPi-SMZR THERMO PLATE F/SMZ DIA BASE R |  | TPi-SMZR Thermo Plate With Controller. Transparent Glass Plate (180mm Diameter). Installs On All SMZ Diascopic Bases (Replaces Standard Glass Insert). | ThermoPlateDia:1:1:P
84455 | TPI-SMZ25X THERMO PLATE F/SMZ25 P2 BASE |  | TPi-SMZ25X Thermo Plate With Controller. Transparent Special Glass Plate (238mm x 251mm) Made Of Super Durable Material To Resist Breakage. Installs On SMZ 18 and SMZ 25 Diascopic Bases (P2-PB, P2-DBL, P2-DBF). | ThermoPlateP2:1:1:P
84456 | TPI-UNIX THERMO PLATE F/ALL DIA BASES |  | TPi-UNIX Thermo Plate With Controller. Transparent Special Glass Plate (211mm W x 430mm L x 75-100mm H) Made Of Super Durable Material To Resist Breakage. For Use With Most Diasciopic Stereo Bases. | ThermoPlateDia:1:1:P
84460 | TPID-SMZ25DX LARGE HEATING PLATE-STEREOS |  | TPiD-SMZ25DX Large Surface Thermo Plate. Transparent Special Glass Plate (248mm x 370mm) Made Of Super Durable Material To Resist Breakage. Installs On SMZ 18 and SMZ 25 Diascopic Bases (P2-PB, P2-DBL, P2-DBF). | ThermoPlateP2:1:1:P
Accessories
83314 | 6V/20W HALOGEN BULB W/REFLECTOR
83326 | 180MM BLACK ANODIZED PLATE F/83311
83361 | FL-150 WHITE BULB /F83360 FL RINGLIGHT
83380 | EKE 21 VOLT 150 WATT BULB FOR C-NI-150
83381 | C-EKZ 10.8V/30W BULB FOR C-NI-30
84246 | 6V-30W BULB
MXA20001 | CLEAR GLASS PLATE 90MM-REPLACE |  |  | 90mmStagePlate:0:1:P
MXA20002 | OPAL/BLK ACRYLIC STG PLATE 90MM-REPLACE |  |  | 90mmStagePlate:0:1:P
MXA22061 | COVER TYPE 104-LAB/OPT/E400/E600/SMZ
MXA25010 | FLUOR LIGHT SHIELD F/SMZ1000/800
MXA25022 | C-GSP2 GLASS PLATE-177MM DIAMETER |  |  | 177mmStagePlate:1:1:P
MXA25023 | C-ASP2 OPAL/BLACK ACRYLIC PLATE-177MM |  |  | 177mmStagePlate:1:1:P
MXA25012 | C-GSP GLASS PLATE 180MM F/NEW DIA STAND |  |  | 177mmStagePlate:1:1:P
MBF92200 | C-AC AC ADAPTER 24 VOLT-5 AMP |  |  | CACAdapter:1:1:P | PowerCord:1:1:H
JBAM-SMZFocusRelay | SMZ Motorized Focus Unit Cable Attachments |  |  | SMZFocusRelay:1:4:P
MNN40920 | Polarizinf Attachment for SMZ |  |  | SMZPolAttachment:1:1:P | SMZOBJAttachment:1:1:P

## SMZ1270¦800
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
Stands
MNA51000 | SMZ800N ZOOM BODY-NC | Base |  | SMZCheapZoomBody:1:1:P | PRNNosepiece:0:1:H | SMZ800Objective¦SMZ1270Objective:1:1:H | SMZ8001270Tube:1:1:H |  | SMZIntermediateTube:0:6:H
MNA52100 | SMZ1270 ZOOM BODY-NC | Base |  | SMZCheapZoomBody:1:1:P | PRNNosepiece:0:1:H | SMZ1270Objective:1:1:H | SMZ8001270Tube:1:1:H |  | SMZIntermediateTube:0:6:H
MNA52110 | SMZ1270I INTELLIGENT ZOOM BODY-NC | Base | Requires 97050 USB cable | SMZCheapZoomBody:1:1:P | PRNiNosepiece:0:1:H | SMZ1270Objective:1:1:H | SMZ8001270Tube:1:1:H |  | SMZIntermediateTube:0:6:H
MMA36400 | SMZ745 CH ZOOM STEREO MICROSCOPE | Base |  | SMZCheapZoomBody:1:1:P | SMZEyepiece:2:2:H | SMZ745Objective:0:1:H | SMZOBJAttachment:0:1:H
MMA36410 | SMZ745T CH ZOOM STEREO MICROSCOPE-TRINOC | Base |  | SMZCheapZoomBody:1:1:P | SMZEyepiece:2:2:H | SMZ745Objective:0:1:H | SMZOBJAttachment:0:1:H | CmountCamera:0:1:H
Bodytubes; Eyepieces
MNB52110 | SMZ1270/800N P-B BINOCULAR TB-NC |  | Compatible with Diascopic Bases, MMD31001, MMD31011, MMD31100 & MMD23500 | SMZ8001270Tube:1:1:P |  | SMZEyepiece:2:2:H
MNB53530 | SMZ1270/800N P-TERG100 TILT TRINOC TB-NC | BlueCircle | NOT compatible with Diascopic Bases, MMD31001, MMD31011, MMD31100 & MMD23500 | SMZ8001270Tube:1:1:P | FmountAdapter:0:1:H | SMZEyepiece:2:2:H
MNB53540 | SMZ1270/800N P-TERG50 TILT TRINOC TB-NC | BlueCircle | NOT compatible with Diascopic Bases, MMD31001, MMD31011, MMD31100 & MMD23500 | SMZ8001270Tube:1:1:P | FmountAdapter:0:1:H | SMZEyepiece:2:2:H
MNB53551 | SMZ1270/800N P-TL100 TRINOC TUBE2-NC |  | Compatible with Diascopic Bases, MMD31001, MMD31011, MMD31100 & MMD23500 | SMZ8001270Tube:1:1:P | FmountAdapter:0:1:H | SMZEyepiece:2:2:H
MXA22132 | CFI REPLACEMENT EYEGUARD F/22 & 25MM-NC
Objectives for SMZ800N
MNH43055 | SMZ-800N P-ACHROMAT 0.5X OBJECTIVE-CH |  | WD=189mm | SMZ800Objective:0:1:P | SMZOBJAttachment:0:1:H
MNH43105 | SMZ-800N 1X PLAN ACHROMAT OBJECTIVE-CH |  | WD=78mm | SMZ800Objective:0:1:P | SMZOBJAttachment:0:1:H
Objectives for SMZ1270/1270i &SMZ 800N
MNH54050 | SMZ1270/800N P-PLAN APO 0.5X/WF |  | WD=82mm | SMZ1270Objective:0:1:P | SMZOBJAttachment:0:1:H
MNH54076 | SMZ1270/800N P-PLAN APO 0.75X/WF CH |  | WD=107mm | SMZ1270Objective:0:1:P | SMZOBJAttachment:0:1:H
MNH54105 | SMZ1270/800N P-PLAN APO 1X/WF CH |  | WD=70mm | SMZ1270Objective:0:1:P | SMZOBJAttachment:0:1:H
MNH54100 | SMZ1270/800N P-PLAN APO 1X/WF |  | WD=70mm | Alias:MNH54105
MNH54150 | SMZ1270/800N P-ED PLAN 1.5X/WF |  | WD=44mm | SMZ1270Objective:0:1:P | SMZOBJAttachment:0:1:H
MNH54200 | SMZ1270/800N P-ED PLAN 2X/WF |  | WD=35mm | SMZ1270Objective:0:1:P | SMZOBJAttachment:0:1:H
Objectives for SMZ745
MMH31055 | G O.5X AUX OBJECTIVE-CH |  | WD=211mm | SMZ745Objective:1:1:P
MMH31075 | G 0.7X AUX OBJECTIVE-CH |  | WD=150mm | SMZ745Objective:1:1:P
MMH31155 | G 1.5X AUX OBJECTIVE-CH |  | WD=61mm | SMZ745Objective:1:1:P
MMH31205 | G 2X AUX OBJECTIVE-CH |  | WD=43.5mm | SMZ745Objective:1:1:P
Nosepieces
MNP51610 | SMZ1270 P-RNI2 INTELLIGNT NOSEPC-NC |  |  | PRNiNosepiece:1:1:P
MNP51620 | SMZ1270/800N P-RN2 NOSEPIECE-NC |  |  | PRNNosepiece:1:1:P
Diascopic Bases (Built-in Focus Mount)
MMD31001 | C-PSN PLAIN STAND / NC | GreenSquare | 292mm (W) x 305mm (D) x 325mm (H) with focus mount & 90mm black & opal plate. Compatible with Tri-noc MNB53551 (only), or use Bi-noc and beamsplitter. Not Compatible with P-EFL Epi Fluorescense Attachment. | ExtensionRingAdapterAchro:0:1:H | 90mmStagePlate:0:1:P | SMZTiltingStage:0:1:H |  |  | SMZCheapZoomBody:1:1:H
MMD31011 | C-PSCN PLAIN STAND / NC | GreenSquare | 180mm (W) x 265mm (D) x 325 (H) with focus mount & 90mm black & opal plate. Not recommended for use with phototube models or camera systems. Compatible with Bi-noc Tube and SMZ-800N or SMZ-1270 only. NOT compatible with P-EFL Epi Fluorescense Attachment. | Alias:MMD31001
MMD23500 | C-LED EPI/DIA STAND W/ FOCUS MOUNT | GreenSquare | Solid & clear glass 90mm plates (See DB08-002). Requires 79035 power cord. Compatible with Tri-noc MNB53551 (only), or use Bi-noc and beamsplitter. Not Compatible with P-EFL Epi Fluorescense Attachment. | ExtensionRingAdapterAchro:0:1:H | 90mmStagePlate:0:1:P | SMZTiltingStage:0:1:H | PowerCord:1:1:H |  | SMZCheapZoomBody:1:1:H
MMD31100 | C-DS STAND | GreenSquare | Stand with tilting mirror. 180mm stage plate. Requires MME42310 LED Illuminator. Compatible with Tri-noc MNB53551 (only), or use Bi-noc and beamsplitter. Not Compatible with P-EFL Epi Fluorescense Attachment. | ExtensionRingAdapterAchro:0:1:H | 177mmStagePlate:0:1:P | SMZSlidingStage:0:1:H | SMZDarkFieldUnit:0:1:H | CDLSU:1:1:H | SMZCheapZoomBody:1:1:H |  |  | ThermoPlateDia:0:1:H
MME42310 | SMZ1270/800N C-DSLU2 LED UNIT F/DIA-NC |  | AC adapter (MQF52057) and power cord (79035) need to be ordered separately. | CDLSU:1:1:P | ACAdapter:1:1:H
Diascopic Bases (Requires Focus Mount)
MMD33010 | C-ER EXTENSION RING ADAPTER FOR 0.5X |  | The 56mm Extension Ring can be used in-lieu-of the Extension Pillar if the customers sample height will not exceed 50mm and the SMZ-800N is combined with the P-Achromat 0.5x Objective (MNH43055). For this combination and sample heights greater than 50mm, use the Extension Pillar (MND43025).
Boom Stands (Requires Focus Mount)
MMD23032 | UNIVERSAL BOOM STAND-3, HEAVY DUTY BASE |  | Heavy duty double arm boom stand with 24.5mm post. Requires MND44002 (C-FMAN) focus mount. | SMZFMAN:1:1:H
MMD33030 | UNIVERSAL BOOM STAND-2 WITH SQUARE BASE |  | Requires MND44011 (C-FMB) focus mount. | SMZFMBN:1:1:H
MMD33530 | UNIVERSAL BOOM STAND-2 WITH SQUARE BASE-CH |  | Requires MND44011 (C-FMB) focus mount. | SMZFMBN:1:1:H
Focus Mounts
MND54000 | C-FMCN FOCUS MOUNT-NC |  | For MNDxxxxx base with 32mm post. | SMZFMCN:1:1:P | SMZCheapZoomBody:1:1:H
MND44011 | C-FMBN BONDER ARM MOUTN BN-NC |  | Bonder arm post diameter is 15.8mm. Not compatible with dual nosepieces (P-RN2 and P-RNI2). | SMZFMBN:1:1:P | SMZCheapZoomBody:1:1:H
MND44002 | C-FMAN FOCUS MOUNT AN-NC |  | Mounts to 24.5mm post. Not compatible with dual nosepieces (P-RN2 and P-RNI2). | SMZFMAN:1:1:P | SMZCheapZoomBody:1:1:H
Teaching Head Accessories
MNB45100 | P-THSS TEACHING HEAD |  | Can not be used silultaneously with MNE5433 | SMZCheapZoomBody:1:1:H | TeachingStand:1:1:H | LEDPointer:0:1:H | SMZ8001270Tube:1:1:H
MBB75145 | Y-THPL LED POINTER UNIT CH |  |  | LEDPointer:1:1:P | ACAdapter:1:1:H
MND45100 | P-THS TEACHING HEAD STAND |  | Stand for MNB45100 Teaching Head. | TeachingStand:1:1:P
Phototube & Intermediate Accessories
MNB45901 | P-BEAM SPLITTER 2 |  | Accepts 0.55x MQD42056 or MBB63435 (LV-TV ISO Adapter). Can not be used silultaneously with MNE5433 | SMZIntermediateTube:1:1:P | FmountAdapter:0:1:H
MNB35200 | P-IRIS DIAPHRAGM F/SMZ800-SMZ10A |  |  | SMZIntermediateTube:1:1:P
MNB45400 | P-IDT DRAWING TUBE |  | Can not be used silultaneously with MNE5433 | SMZIntermediateTube:1:1:P
MXA25000 | P-IER EYELEVEL RISER |  | Maximum of three without intermediate modules or phototube | SMZIntermediateTube:1:1:P
Fluorescence Illumination
MNE54332 | SMZ1270/800N P-EFL2 EPI FLUOR ATTACHMNT | BlueCircle | Can not be used simultaneously with any other intermediate tubes. NOT compatible with any of the Diascopic Bases with the built-in-focus. NOTE: Accepts direct attachment of the D-LEDI Fluorescence LED Illumination System and also comes with a fiber adpater for applications requiring a Fiber Illuminator. The D-LEDI LED Fluorescence Illumination System is sold separately. | SMZIntermediateTube:1:1:P | DLED¦LLG:1:1:H | SMZ1270800EpiCube:1:4:H
Fluorescence cubes
MXA25021 | SMZ1270/800N P-EFLBF FILTER CUBE-EPI BF |  |  | SMZ1270800EpiCube:1:1:P
83420 | SMZ1270/800N P-EFLC BFP AT FILTER SET |  | EX 395/25, DM 415, EM 460/50 | SMZ1270800EpiCube:1:1:P
83421 | SMZ1270/800N P-EFLC ECFP AT FILTER SET |  | EX 435/20, DM 455, EM 480/30 | SMZ1270800EpiCube:1:1:P
83422 | SMZ1270/800N P-EFLC GFP LP AT FILTER SET |  | EX 480/30, DM 505, EM 515 LP | SMZ1270800EpiCube:1:1:P
83423 | SMZ1270/800N P-EFLC GFP BP AT FILTER SET |  | EX 480/30, DM 505, EM 535/40 | SMZ1270800EpiCube:1:1:P
83424 | SMZ1270/800N P-EFLC EYFP AT FILTER SET |  | EX 495/20, DM 515, EM 540/30 | SMZ1270800EpiCube:1:1:P
83425 | SMZ1270/800N P-EFLC TRITC AT FILTER SET |  | EX 540/25, DM 565, EM 605/55 | SMZ1270800EpiCube:1:1:P
77074448 | 39010 TR/MC/AXFL594 SMZ1270/800N P-EFLC |  |  | SMZ1270800EpiCube:1:1:P
77074976 | 49006 ET CY5 SET MTED SMZ1270/800N CUBE |  |  | SMZ1270800EpiCube:1:1:P
77074943 | 59004 ET FITC/TRITC MTD SMZ1270/800N CB |  | 59004 ET FITC/TRITC Filter Set; mounted SMZ1270/800N cube. (2 emitters) | SMZ1270800EpiCube:1:1:P
Epi Components
MME42210 | SMZ1270/800N C-LSL2 LED EPI ILLUM-NC |  | Can be used on MMD31011, MMD31001 and MMD31100 Stands. AC adapter (MQF52056) and power cord (79035) need to be ordered separately. | CLSL2:1:1:P
MME32210 | G-EIA FLEXIBLE GOOSENECK ILLUMINATOR ARM |  | Can be used on MMD31011, MMD31001 and MMD31100 Stands. | GEIA:1:1:P
MNB55820 | SMZ1270/800N P-CI COAXIAL EPI ILLUM |  |  | SMZIntermediateTube:1:1:P | CFLED:1:1:P
MNB55821 | SMZ1270/800N P-CI QL1X 1/4 LAMBDA PLAT |  | For 0.5x, 0.75x and 1.0x objectives.
MNB55822 | SMZ1270/800N P-CI QL1.5X 1/4 LAMBDA PLAT |  | For 1.5x objective.
MXK60559 | LED RING ILLUM SM-LW61J15 100-240V-NC |  | SM-LWD61J15 LED Ring Illuminator 100-240V. White LED ring illuminator with a mount diameter of 60 mm. Comes with a brightness control dial. An AC adapter (MQF52057) and power cord (79035) need to be ordered separately. |  | ACAdapter:1:1:H
MXK60515 | LED RING ILLUM SM-LW61J14 100-240V-NC |  |  | Alias:MXK60559
83370 | DUAL GOOSENCK W/ FOC LEN ATTACH TO STAND |  | Attaches to MMD31000, MMD31010, MMD31100. Compatible with 83364, 83365 and 83359 Illuminators.
83369 | C-ADPTR & FBR OPTIC CBLE F/C-DS DIA STND |  | Single light guide with adapter for C-DS Diascopic base MMD31100. Compatible with 83364, 83365 and 83359 Illuminators.
83366 | RINGLIGHT C FOR SMZ 445/800/1000/1500 |  | Fits all stereo objectives except ED Plan 2x WF. Compatible with 83364, 83365 and 83359 Illuminators.
83367 | POLARIZER FOR RINGLIGHT C
83360 | C-FLUOR RING LIGHT
Thermal Stages
84434 | TP-SMZR THERMO PLATE FOR SMZ DIA BASE R |  | Tokai Hit - demo stock Only
84435 | TP-SMZ25 THERMO PLATE SMZ18/25 P2 BASES |  | Tokai Hit - demo stock only

## SMZ445¦745
Product Code | Description | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
MMA23008 | SMZ-445 ZOOM STEREO MICROSCOPE-NC
MMA23007 | SMZ-460 ZOOM STEREO MICROSCOPE-NC
MMK20106 | E10XB-SMZ 445/460 EYEPIECE-NC | Economy Type for SMZ-445 - Does Not Accept Reticle
MMK20152 | E15XB-SMZ 445/460 EYEPIECE-NC | Economy Type for SMZ-445 - Does Not Accept Reticle
MMK20208 | E20XB-SMZ 445/460 EYEPIECE-NC | Economy Type for SMZ-445 - Does Not Accept Reticle
MXA25019 | M SM E 10X EYEPIECE RUBBER RING / NC
MMH23005 | 0.5X AUX OBJ-CH F/SMZ-445/460/1/1B/2B/2T | WD=181mm
MMH23007 | 0.7X AUX OBJ-CH F/ SMZ-445/460/1/1B/2B/2 | WD=127.5mm
MXA20563 | STEREO STAGE CLIP
MME32000 | G-ICIL COAXIAL EPISCOPIC LED ILLUMINATOR | See DB10-037
MME39200 | G-0BA RINGLIGHT ADAPTER
83374 | C-RINGLIGHT A, F/SMZ-745/645/660 (ONLY) | Compatible with 83364, 83365 and 83359 Illuminators.

## LAPP
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
91170 | 50/50 MIRROR FOR L-APP ILLUMINATOR |  | Main Branch Mirror - Required | SwitchingMirror:1:1:P
91171 | 488NM REFLECT/561NM TRANSMIT DUAL TIRF |  | Main Branch Mirror - Required | SwitchingMirror:1:1:P
91172 | 405NM REFLECT/430-800NM TRANS DUAL TIRF |  | Main Branch Mirror - Required | SwitchingMirror:1:1:P
91173 | 405NM TRANSMIT/REFLECT 320NM & 720NM |  | Main Branch Mirror - Required | SwitchingMirror:1:1:P
99224 | 405NM STIM DICHROIC MIRROR F/TI2 LFOV V2
77015955 | SUPPORT STAND F/PHOTOACT PT SCANNER |  | Mounting Stand
77016061 | SHORT GANTRY STAND |  | Mounting Stand
77064016 | POLYGON1000 DYNAMIC SPATIAL ILLUMINATOR | NIS-POLYGON? | 1x Front Lens Unit. 11.6mm diagonal FOV, 7.6um pixel size. | Polygon1000:1:1:P | Polygon10002XLens:0:1:H | LLG:1:1:H
77064018 | POLYGON1000 OPT ATTACHMENT F/LAPP ILLUM | NIS-POLYGON? | Connects to main or sub branch of lapp | LAPPModuleU:1:1:P | Polygon1000:1:1:H
77064019 | POLYGON1000 DMD W/2X FRONT LENS | NIS-POLYGON? | 2x Front Lens Unit. 23.2mm diagonal FOV, 15.2um pixel size. | Polygon1000:1:1:P |  | LLG:1:1:H
77064017 | FRONT TUBE OPTICAL ASEMBLY F/POLYGON1000 |  | 2x front lens a la carte (can replace the 1x lens if a user wants both options interchangeable – sell with 77064016 as an additional option) | Polygon10002XLens:1:1:P
77086005 | PHOTOSIM SCANNER W/OPTICS F/TI2 LAPP | NIS-XYGALVO?,Triggering |  | LAPPModuleU:1:1:P | A1out¦A2out¦A3out¦A4out:1:1:H | BNC:2:2:H
77086009 | UV OMS WITH LASER SAFETY HARDWARE
MEB58100 | Ti-BPU BACK PORT UNIT |  |  | LAPPU:1:1:P
MED53210 | TI2-LA-SUE TI2-E STAGE UP KIT | StageUp | A Dual Stack System requires a stage up kit and second turret | StageUpE:1:1:P | EpiTurretE¦EpiTurretA:0:1:H | LAPPE¦LAPPA¦LAPPU:0:1:H | StageUpLens:0:1:H
MED53220 | TI2-LA-SU TI2-A/U STAGE UP KIT | StageUp | A Dual Stack System requires a stage up kit and second turret | StageUpU:1:1:P | EpiTurretA:0:1:H | LAPPA¦LAPPU:0:1:H | StageUpLens:0:1:H
MED55200 | TI-LA-EXTENSION SUPPORT SPACER |  | Large LApp modules on the upper stack require an extenstion to their support leg
MEE54730 | TI2-LA-FLL EPI FL MODULE L FOV |  |  | LAPPModuleU:1:1:P | LLG:1:1:H
MEE54740 | TI2-LA-FL-2 EPI-FLUOR MODULE |  |  | LAPPModuleU:1:1:P | FIBA¦DLED:1:1:H
MEE54840 | TI2-LA-BM-E MOTORIZED MAIN BRANCH | MainBranch | Ti2-E Only - Connects to Ti2 Control Box | LAPPE:1:1:P | LAPPModuleU¦LAPPModuleE:1:1:H | LAPPSubBranchU¦LAPPSubBranchE:0:1:H | SwitchingMirror:1:1:H
MEE54850 | TI2-LA-BM MAIN BRANCH | MainBranch |  | LAPPE:1:1:P | LAPPModuleU¦LAPPModuleE:1:1:H | LAPPSubBranchU:0:1:H | SwitchingMirror:1:1:H
MEE54860 | TI2-LA-BF FIXED MAIN BRANCH | MainBranch |  | LAPPU:1:1:P | LAPPModuleU:1:1:H
MEE54870 | TI2-LA-BS-E MOTORIZED SUB BRANCH |  | Ti2-E Only -Can use only 1 per stand | LAPPSubBranchE:0:1:P | LAPPModuleU¦LAPPModuleE:0:2:H
MEE54880 | TI2-LA-BS SUB BRANCH |  |  | LAPPSubBranchU¦LAPPSubBranchE:0:1:P | LAPPModuleU:1:1:H
MEE58017 | TI-LA-EXCHAGEABLE MIRROR (100% REFLECT) |  | Main Branch Mirror - Required | SwitchingMirror:1:1:P
MEE58020 | TI-LA-FRAP FRAP MODULE |  |  | LAPPModuleU:1:1:P |  |  | A1out¦A2out¦A3out¦A4out:1:1:H
MEE59130 | TI-LA-TIRF TIRF MODULE | TIRF |  | LAPPModuleU:1:1:P |  |  | A1out¦A2out¦A3out¦A4out:1:1:H
MEE59210 | TI-LA-HTIRF H-TIRF MODULE | TIRF |  | LAPPModuleU:1:1:P | HTIRFController:1:1:H |  | A1out¦A2out¦A3out¦A4out:1:1:H
MEE59251 | TI2-LA-NS2 N-STORM MODULE 2 | TIRF |  | LAPPModuleE:1:1:P | HTIRFController:1:1:H |  | A1out¦A2out¦A3out¦A4out:1:1:H
MEE59250 | TI-LA-NS N-STORM MODULE | TIRF |  | Alias:MEE59251
MEE59220 | TI2-LA-TIRF-E MOTORIZED TIRF MODULE | TIRF |  | LAPPModuleE:1:1:P | A1out¦A2out¦A3out¦A4out:1:1:H | NSTORMMag:0:1:H
MEE54330 | TI2-NS2-MU N-STORM MAG CHANGE UNT |  |  | NSTORMMag:1:1:P
MED54301 | TI-NSM N-STORM KIT |  |  | SidePortAdapter:1:1:P | CmountAdapter:1:1:H
MED54500 | TI-NS2 N-STORM KIT |  |  | SidePortAdapter:1:1:P | CmountAdapter:1:1:H
MEF55805 | TI2-LA-CTR H-TIRF CONTROLLER |  |  | HTIRFController:1:1:P
MEV53001 | TI-LA-LAD LASER ADAPTER |  | Light guide adapter - Required
MEV53002 | TI-LA-LEDAD LED ADAPTER |  | Light guide adapter - Required
MEV53004 | TI-LA-SUL STAGE UP LENS | StageUpLens |  | StageUpLens:1:1:P
MXA22198 | D-LEDI RETROFIT ADAPTER
MXA22205 | C-FIBA ADAPTER FOR FIBER LIGHT |  |  | FIBA:1:1:P | LLG:1:1:H
MEE54741 | TI2-LA-FL-3 EPI-FLUOR MODULE |  | Improved NIR excitation transmission | Alias:MEE54740 | 25mmFilter:0:2:H
MXA22197 | C-FIBA ADAPTER FOR FIBER LIGHT |  |  | Alias:MXA22205

## Spinning Disk
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
W1 and ZIVA Unit Bundles
94030 | CSU-W1 50UM PH SINGLE DISK 1 CAM W/ZIVA | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H | ZIVALaser:1:1:H |  | RS232Adapter:1:1:H | W1Dichroic:0:3:H | AddPart:JBAM-ZIVA,77019802 | W1SingleLayerLens:0:1:H
94032 | CSUW1 50UM PH SNGE DSK 2 CAM SYS W/ZIVA | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | ZIVALaser:1:1:H | W1Splitter:1:3:H | RS232Adapter:1:1:H | W1Dichroic:0:3:H | AddPart:JBAM-ZIVA,77019802 | W1SingleLayerLens:0:1:H
94034 | CSU-W1 50/25 DUAL DISK 1 CAM SYS W/ZIVA | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H | ZIVALaser:1:1:H |  | RS232Adapter:1:1:H | W1Dichroic:0:3:H | AddPart:JBAM-ZIVA,77019802 | W1SingleLayerLens:0:1:H
94035 | CSU-W1 50/25 DUAL DISK 2 CAM SYS W/ZIVA | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | ZIVALaser:1:1:H | W1Splitter:1:3:H | RS232Adapter:1:1:H | W1Dichroic:0:3:H | AddPart:JBAM-ZIVA,77019802 | W1SingleLayerLens:0:1:H
94031 | CSU-W1 DL DSK SNGL CAM 50/50UMSR W/ZIVA | Scanhead,NIS-W1?,NIS-3D-DECON? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H | ZIVALaser:1:1:H |  | RS232Adapter:1:1:H | W1Dichroic:0:3:H | AddPart:JBAM-ZIVA,77019802 | W1SingleLayerLens:0:1:H
94033 | CSUW1 DL DSK/CAM SYS 50/50UM SR W/ZIVA | Scanhead,NIS-W1?,NIS-3D-DECON? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | ZIVALaser:1:1:H | W1Splitter:1:3:H | RS232Adapter:1:1:H | W1Dichroic:0:3:H | AddPart:JBAM-ZIVA,77019802 | W1SingleLayerLens:0:1:H
JBAM-ZIVA | Lumencor Ziva | LaserSafety,NIS-WLSWITCHER? |  | ZIVALaser:1:1:P | InterlockPhone:1:1:H
W1 and Laser Unit Bundles
99927 | CSU-W1 50UM PH SINGLE DISK 1 CAM W/LASER | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H |  | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H | AddPart:77097048,77019802 | W1SingleLayerLens:0:1:H
99928 | CSU-W1 50UM PH SINGLE DISK 2 CAM W/LASER | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | W1Splitter:1:3:H | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H | AddPart:77097048,77019802 | W1SingleLayerLens:0:1:H
99929 | CSU-W1 DUAL DISK 1 CAMERA SYS W/LASER | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H |  | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H | AddPart:77097048,77019802 | W1SingleLayerLens:0:1:H
99930 | CSU-W1 DUAL DISK 2 CAMERA SYS W/LASER | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | W1Splitter:1:3:H | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H | AddPart:77097048,77019802 | W1SingleLayerLens:0:1:H
99136 | CSU-W1 DL DSK SNGL CAM 50/50UMSR LSR SYS | Scanhead,NIS-W1?,NIS-3D-DECON? | This is CSU-W1 SoRa and mag changer laser bundle. 60" table recommended, 48" minimum (camera will extend beyond table). | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H |  | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H | AddPart:77097048,77019802 | W1SingleLayerLens:0:1:H
99186 | CSU-W1 DL DSK DL CM 50UM/50UM SR LSR SYS | Scanhead,NIS-W1?,NIS-3D-DECON? | This is CSU-W1 SoRa and mag changer laser bundle. 60" table recommended, 48" minimum (camera will extend beyond table). | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | W1Splitter:1:3:H | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H | AddPart:77097048,77019802 | W1SingleLayerLens:0:1:H
W1 head only
99923 | CSU-W1 50UM PH SINGLE DISK 1 CAMERA SYS | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H |  | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H |  | W1SingleLayerLens:0:1:H
99924 | CSU-W1 50UM PH SINGLE DISK 2 CAMERA SYS | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | W1Splitter:1:3:H | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H |  | W1SingleLayerLens:0:1:H
99925 | CSU-W1 DUAL DISK 1 CAMERA SYSTEM | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H |  | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H |  | W1SingleLayerLens:0:1:H
99926 | CSU-W1 DUAL DISK, 2 CAMERA SYSTEM | Scanhead,NIS-W1? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | W1Splitter:1:3:H | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H |  | W1SingleLayerLens:0:1:H
99129 | CSU-W1 DUAL DISK SINGLE CAM 50UM/50UM SR | Scanhead,NIS-W1?,NIS-3D-DECON? |  | SidePortAdapter:1:1:P | CmountCamera:1:1:H | 25mmFilter:0:10:H |  | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H |  | W1SingleLayerLens:0:1:H
99158 | CSU-W1 DUAL DISK DUAL CAM 50UM/50UM SR | Scanhead,NIS-W1?,NIS-3D-DECON? |  | SidePortAdapter:1:1:P | CmountCamera:1:2:H | 25mmFilter:0:20:H | W1Splitter:1:3:H | W1Dichroic:0:3:H | RS232Adapter:1:1:H | P1out¦P2out¦P3out¦P4out:1:1:H |  | W1SingleLayerLens:0:1:H
W1 Accessories
99298 | CSU-W1 ADPTR F/TI2 SNGL LAYER 0.95X MAG |  | Optional adapter for CSU-W1 when used on single layer non-stageup configuration of Ti2. Introduces 0.95x magnification factor. | W1SingleLayerLens:1:1:P
99933 | DM A561LP |  | one beamsplitter is required for dual camera models only (#99924 or #99926) | W1Splitter:1:1:P
99934 | DM A514LP |  | one beamsplitter is required for dual camera models only (#99924 or #99926) | W1Splitter:1:1:P
99935 | DUMMY GLASS |  | one beamsplitter is required for dual camera models only (#99924 or #99926) | W1Splitter:1:1:P
99936 | FULL MIRROR |  | one beamsplitter is required for dual camera models only (#99924 or #99926) | W1Splitter:1:1:P
99937 | DM 405/488/561 |  |  | W1Dichroic:1:1:P
99938 | DM 445/514/640 |  |  | W1Dichroic:1:1:P
99226 | 445/514/594 CSU W1 DICHROIC 22.2X25.0 MM |  |  | W1Dichroic:1:1:P
99229 | 405/488/594 CSU W1 DICHROIC 22.2X25.0 MM |  |  | W1Dichroic:1:1:P
99237 | 445/515/561 CSU W1 DICHROIC 22.2X25.0 MM |  |  | W1Dichroic:1:1:P
77019761 | 4-PORT RS-232 SERIAL ADAPTER PCIE CARD. |  | Required for PCs without a serial port including all Z4 workstations. A serial to usb adapter could be used instead. | RS232Adapter:0:1:P | PCIEport:1:1:P
77019802 | PCI EXP RS232 SERIAL CNTRLR CARD 1 PORT |  |  | RS232Adapter:0:1:P | PCIEport:1:1:P
77155115 | USB/SERIAL (RS232) CONVERTER |  |  | RS232Adapter:0:1:P
Cerest V3 Packages
99287 | CREST X-LIGHT V3 LFOV25 SPN DSK CONFOCAL | Scanhead,NIS-CREST? | Crest X-Light V3 LFOV25 Spinning Disk Confocal. Includes 7-line solidstate light source, VBCTGRN with despeckling, ~800 mW per line. Dual andpenta band dichroics, excitation, and emission filters included. | AddPart:77059524,77060140,MXA22158,77059528,77059529
99289 | CREST X-LIGHT V3 DEEPSIM SP DSK CONFOCAL | Scanhead,NIS-DEEPSIM?,NIS-CREST? | Hardware package for Crest X-Light V3 Spinning Disk Confocal for Ti2. Includes Celesta 7-line solidstate light source, 1W per channel. V3 scanhead features built-in illuminator, optimized 50/250um pinhole disk, dual F-mount and C-mount adapters for cameras, adapter for Ti2 microscope, motorized 3-position dichroic switcher, motorized 8-position emission filter wheel, and motorized 4-position excitation wheel. Comes with USB cable, power supply, filter holders, filter tools, and Ti2 interlock cable. | AddPart:77059524,77060140,MXA22158,77059528,77059529
99293 | CREST DEEPSIM STAND ALONE SYS | Scanhead,NIS-DEEPSIM? | Hardware package for DeepSIM Standalone (V3 not included and cannot be "upgraded" to V3). Includes Celesta 7-line solidstate light source, 1W per channel. C-mount and adapter for Ti2 microscope, motorized 8-position emission filter wheel. Comes with USB cable, power supply, filter holders, filter tools, and Ti2 interlock cable. | AddPart:77059524,77060140,MXA22158,77059528,77059529
Crest X-Light V3
77059524 | X-LIGHT V3 50-250 FOV25 DS |  |  | SidePortAdapter:1:1:P | CelestaLightSource:1:1:H | CmountCamera¦FmountCamera¦TmountCameraH:1:1:H | CmountCamera¦FmountCamera¦TmountCameraV:0:1:H | 25mmFilter:0:8:H
Accessories
77059513 | 50/500 DISK BOX F/X-LIGHT V2 LFOV |  | Optional separate disk box with wider pinhole spacing for very thick samples.
77059528 | HORIZONTAL TELEDYNE CAMERA MOUNT FOR V3 |  | *now included in all V3 boms, provides better stability for Photometrics cameras | TmountCameraH:1:1:P | TmountCamera:1:1:H
77059529 | VERTICAL TELEDYNE CAMERA MOUNT FOR V3 |  | *now included in all V3 boms, , provides better stability for Photometrics cameras | TmountCameraV:1:1:P | TmountCamera:1:1:H
77059527 | CREST OPTICS DICHORIC HOLDER ASSEMBLY FOR V3 |  | Extra holder if adding another dichroic
Crest X-Light V2 LFOV25 Celesta Package
99286 | CREST X-LIGHT V2 LFOV25 SPN DSK CONFOCAL | Scanhead,NIS-CREST? | Updated version, despeckled Celesta
99259 | CREST X-LIGHT V2 LFOV25 SPN DSK CONFOCAL | Scanhead,NIS-CREST? | Previous version, non-despeckled Celesta
Other accessories
77012923 | LA-C-NF NIKON F TO C MOUNT LENS ADAPTER
Crest Cicero Packages
99899 | CICERO, 50/250UM WITH CELESTA 5CH | Scanhead | Laser System! Only use with Ti2/TiE/Ni/FN1 with Laser-Safety Components | SidePortAdapter:1:1:P | CelestaLightSource:1:1:H | AddPart:JBAM-CELESTA5CH
99898 | CICERO, 60/220UM WITH AURA 4CH | Scanhead | LED System, no laser safety components required | SidePortAdapter:1:1:P | LLG:1:1:H | AddPart:JBAM-AURA4CH
JBAM-CELESTA5CH | Celesta 5 Channel | LaserSafety,NIS-WLSWITCHER? |  | CelestaLightSource:1:1:P | InterlockPhone:1:1:H
JBAM-AURA4CH | Aura 4 Channel | NIS-WLSWITCHER? |  | LumencorLLGSource:1:1:P
Additional Optional Adapters
77059535 | LFOV MT ADAPTER FROM NI-E LFOV QUADROC |  | Input adapter for quadrocular head
77059534 | X-LIGHT TO MICROSCOPE C-MOUNT
Additional Optional Emission Filters

## 3rd Party
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
XY Stages
For Ti/Ti2 Inverted Microscopes (stage inserts required)
77011331 | PROSCAN II MTRZD FLT STG 4X3 Encoded | NIS-3PSTAGES? | H117 Flat topped stage, 4x3" travel, 50 nm min step size and encoded resolution | StageH117:1:1:P | HolderH117:1:6:H | PriorController:1:1:H | PriorJoystick:1:1:H
77011329 | PROSCAN II MTRZD FLT STG 4X3, 10NM RESOL | NIS-3PSTAGES? | H117 Flat topped stage, 4x3" travel, nonencoded, 10 nm min step size | StageH117:1:1:P | HolderH117:1:6:H | PriorController:1:1:H | PriorJoystick:1:1:H
77011446 | SLIDE HOLDER, 1X3 INCH SLIDE F/H117 STG |  | Fit Prior H117 stages (small opening), no piezo stage, extra recessed low profile, Prior part H224XRLP. | HolderH117:1:1:P
77011401 | 4 SLIDE HOLDER F/H117 STG |  | Fits Prior H117 stages (small opening), no piezo stage, Prior part H234XR. | HolderH117:1:1:P
77011671 | FLASK HOLDER (T25 FLASK) FOR H117 STAGE |  |  | HolderH117:1:1:P
77011707 | HOLDER F/T225 CORNING FLSK |  | Sample holder bar kit, T225 cornign flask #431082, fits H117 stage top.Holds T225 flask on stage top surface. May require objective spacers forachieving proper focus. | HolderH117:1:1:P
77011672 | FLASK HOLDER (T75 FLASK) FOR H117 STAGE |  |  | HolderH117:1:1:P
77011422 | PETRI DISH HLDR VAR DIAMETER F/H117 |  | Fits Prior H117 stages (small opening), no piezo stage, variable diameter petri dishes 25-86mm, Prior part H229XR. | HolderH117:1:1:P
For Ni UPright Microscopes (with fixed nosepieces), AZ100, or SMZ18/25 Microscopes (stage inserts required)
77011510 | H101F PS III MTZD UP STG 4X3, 10NM RES | NIS-3PSTAGES? | H101F Flat topped stage, 4x3" travel, encoded for 0.1um resolution | StageH101:1:1:P | HolderH101:1:6:H | PriorController:1:1:H | PriorJoystick:1:1:H
77011509 | H101F PS III MTZD UP STG 4X3, NO SCALES | NIS-3PSTAGES? | H101F Flat topped stage, 4x3" travel, nonencoded | Alias:77011510
77011522 | SMZ18/25 STG ADPTR F/H101F/H101A/OPTISCN |  | Include this stage adapter for SMZ18/25. | SMZH101Adapter:1:1:P | StageH101:1:1:H
77011511 | H101F SLIDE HOLDER 1X3, 2X3 SLIDES |  | Fits Prior H101F stage (small opening), no piezo, low profile for 3"x1" or 3"x2" slides, Prior part H224ARMF. | HolderH101:1:1:P
77011018 | SLIDE HOLDER 1X3 INCH SLIDE F/H101 STG |  | Single slide holder for 1 x 3 inch slide for upright microscope, for usewith H101 stage. | HolderH101:1:1:P
77011512 | H101F SLIDE HOLDER F/FOUR 1X3 SLIDES |  | Fits Prior H101F stage (small opening), no piezo, holds four 3"x1" slides, Prior part H234F. | HolderH101:1:1:P
77011418 | FOUR SLIDE HLDR RECD FOR H101 STAGE |  | Stage Insert for H101 Stage: Four Slide Holder, Recessed, (Holds Four I x 3 inch Slides) for Upright Stage (4x3 inch travel),, | HolderH101:1:1:P
77011542 | GLASS PLATE FOR PRIOR H101F FLT TP STG |  | Fits Prior H101F stage (small opening), no piezo, Prior part H220F. | HolderH101:1:1:P
77011453 | PSIII UPRT MTRZ STG SYS W/ JSTK & CNTRLR | NIS-3PSTAGES? | ProScan III System: Upright Motorized Microscope Stage. Includes H101ANNN, motorized stage (4x3 inches); H31XYZE controller; PS3J100 joystick. | MISC:0:1:P | AddPart:77011509,77011404,77011688
77011021 | PROSCAN UPRIGHT MOTORIZED STAGE | NIS-3PSTAGES? |  | StageH101:1:1:P | HolderH101:1:6:H | PriorController:1:1:H | PriorJoystick:1:1:H
All stages above require a controller, joystick, and NIS-Elements hardware module for software control
77011404 | PSIII CONTROLLER W/XYZ DRVS, ENCODED |  | Prior H117 and H101F required a Proscan III controller and joystick | PriorController:1:1:P
77011688 | JOYSTICK CONTROL UNIT F/XY CONTROL ONLY |  | New, updated version joystick for XY | PriorJoystick:1:1:P
For Ni and FN1 Upright Microscopes (with focusing nosepieces)
77011469 | Z-DECK PLATFM ENCD F/FN1 INCL STG/CTRL | NIS-3PSTAGES? | Z Deck II flat topped motorized stage (encoded), includes controller, joystick, connection cables, and stands. | NIZDECK:1:1:P
77011694 | Z DECK DUAL Z PLATFORM SYSTEM | NIS-3PSTAGES?
Z Focus drives for manual microscopes
77011383 | NIKON REMOTE FOCUS KIT W/ CONTROLLER |  | Prior NIKRFK, requires microscope adapter.
77011386 | FN1 SCOPE ADAPTER FOR NIKRFK |  | Prior part NIKFN1MA.
77011389 | FN1 ENCODER ADAPTER KIT FOR NIKRFK |  | Optional for FN1, requires scope adapter, Prior part NIKFN1EK.
77011461 | NI-U SCOPE ADAPTER FOR NIKRFK |  | Prior part NIKNIMA.
77011639 | NIKRFK ADAPTER FOR TI2 SCOPE |  | Includes parts for Ti2, Prior part NIKTI2MA.
77011384 | TI SCOPE ADAPTER FOR NIKRFK |  | Includes parts for Ti, Prior part #NIKTIMA
77011385 | 80I SCOPE ADAPTER FOR NIKRFK |  | Prior part NIK80IEK.
77011690 | DIGIPOT CONTROL UNIT FOR Z CONTROL. |  | New, updated version joystick for Z
Stage Top Piezo Z Focus drives
77011674 | GSP450 450UM Z PIEZO STG KIT | NIS-PIEZO?,Triggering | Dual analog input stage-top pizeo w/450um travel for use with Ti2 motorized stages. Includes Universal Insert. Requires specific incubators (see incubator sheets). Preferred choice for A1 driven systems. | StageEInsert:1:1:P | QueensgateHolder:1:6:H | AddPart:77011709 | TokaiIncubatorQueensgate:0:1:H | BNC:1:1:H
77011682 | QGNPS083-NI 600UM Z PIEZO STG KIT | NIS-PIEZO?,Triggering | Dual analog input stage-top pizeo w/600um travel for use with Ti2 motorized stages. Includes Universal Insert. Requires specific incubators (see incubator sheets). Preferred choice for Ax driven systems. | StageEInsert:1:1:P | QueensgateHolder:1:6:H | AddPart:77011709 | TokaiIncubatorQueensgate:0:1:H | BNC:1:1:H
77011709 | UNIV HLDR EXT RECESS F/QUEENSGATE PIEZO |  | Universal Insert for Queensgate 450/600 piezo (extra) | QueensgateHolder:1:1:P
77011678 | GSP450 WELL PLATE HOLDER |  | Well Plate Holder for Prior 450um Stage-top Piezo | QueensgateHolder:1:1:P
77038033 | PI NANO K PLATE SCAN SYS 200UM W/CNTRLR | NIS-PIEZO?,Triggering | Stagetop pizeo w/200um travel for use with Ti and Ti2 motorized stages MEC56100/MEC56120/MEC56130. | StageEInsert:1:1:P | NanoHolder:1:6:H | TokaiIncubatorMCL:0:1:H |  | BNC:1:1:H
77038024 | PI NANO Z UNIV HLDER F/SLD AND DISHES |  | For PINano Z K Plate piezo stage, 110x160 universal holder for slides and dishes, PI part P-736.UHN. | NanoHolder:1:1:P
77038006 | PI NANO Z SLIDE HOLDER |  | For PINano Z K Plate piezo stage, 110x160 for 3"x1" slides, PI part P-736.SHN. | NanoHolder:1:1:P
77038007 | PI NANO Z 200UM PIEZO STG WELL PLT HOLDR |  | For PINano Z K Plate piezo stage, 110x160 for well plates, PI part P-736.WPN. | NanoHolder:1:1:P
77038023 | PI NANO Z 35MM PETRI DISH HOLDER |  | For PINano Z K Plate piezo stage, 110x160 for 35mm petri dish, PI part P-736.PDN. | NanoHolder:1:1:P
91989 | Ti/Ti2 PIEZO STAGE INSERT W/ USB CNTRLR | NIS-PIEZO?,Triggering | MadCityLabs 100um Stagetop piezo for use with Ti and Ti2 motorized stages MEC56100/MEC56120/MEC56130. | StageEInsert:1:1:P | MCLHolder:1:6:H | TokaiIncubatorMCL:0:1:H |  | BNC:1:1:H
77046519 | UNIV SAMPL HLDR NANO-Z-N/N2 PIEZO INSERT |  | For MCL Z100 piezo stage, universal holder for 3"x1" slides and dishes. | MCLHolder:1:1:P
91995 | PETRI DISH HOLDER 35MM FOR PIEZO STAGE |  | For MCL Z100 piezo stage, insert for 35mm dishes, MCL part PDH35-N. | MCLHolder:1:1:P
91994 | MULTI WELL PLATE HOLDER FOR PIEZO STAGE |  | For MCL Z100 piezo stage, insert for multiwell plates, MCL part MPH-N. | MCLHolder:1:1:P
91997 | UNIVERSAL HOLDER FOR PIEZO STAGE |  | For MCL Z100 piezo stage | MCLHolder:1:1:P
91992 | COVER SLIP HOLDER 1X1 FOR 91991 |  | For MCL Z100 piezo stage | MCLHolder:1:1:P
91991 | SLIDE HOLDER 1X3 FOR PIEZO STAGE |  | For MCL Z100 piezo stage | MCLHolder:1:1:P
91998 | MULTIPLE SLIDE HOLDER FOR PIEZO STAGE |  | For MCL Z100 piezo stage, custom insert for 4x slides, maximizes imaging area | MCLHolder:1:1:P
77046518 | Z PIEZO NANOPO SYS 500 MICRON W/CNTRLR | NIS-PIEZO?,Triggering | Closed Loop Z-Axis Piezo Nanopositioning System with 500 Microns of Travel. Includes Nano-Drive Controller. | StageEInsert:1:1:P | MCLHolder:1:6:H | TokaiIncubatorMCL500:0:1:H |  | BNC:1:1:H
77046517 | Z PIEZO NANOPO SYS 200 UM W/CNTRLR F/TI2 | NIS-PIEZO?,Triggering | 200 micron travel with 0.4 nanometer resolution | StageEInsert:0:1:P | MCLHolder:1:6:H | TokaiIncubatorMCL:0:1:H |  | BNC:1:1:H
Tokai Hit Incubators
77025167 | WSKMX/KM/ZILCS ADPTR F/TI2 XY MTRZ STG |  |  | StageEInsert:1:1:P | TokaiIncubator:1:1:H
77025203 | STXG INCB F/WLPLT/DSH/SLD F/FLT MTZ XY |  | For 100% CO2. Requires 77025167 adapter. Cannot upgrade to piezo later. | TokaiIncubator:1:1:P
77025183 | STXF INCB F/WLPLT/DSH/SLD F/FLT MTZ XY |  | For premixed gas. Requires 77025167 adapter. Cannot upgrade to piezo later. | TokaiIncubator:1:1:P
77025166 | TIZ-WX/BX/W/B ADPTR F/TI2 XY MTRZ STG |  |  | StageEInsert:1:1:P | TokaiIncubatorMCL¦TokaiIncubatorQueensgate:1:1:H
77025198 | STXG INCB F/WELP/DSH/SLD F/TI/TI2 PZ STG |  |  | TokaiIncubatorMCL:1:1:P
77025192 | STXF INCB F/WLPLT/DSH/SLD F/TI/TI2 PIEZO |  |  | TokaiIncubatorMCL:1:1:P
77025235 | TIZWX INCB SYS F/HYPOXIA EXPERIMENTS |  |  | TokaiIncubatorMCL:1:1:P
77025181 | STXG INCB SYS F/DSH/SLD F/TI/TI2 PZ STG |  |  | TokaiIncubatorMCL:1:1:P
77025263 | STXF INCB SYS F/DSH/SLD F/TI/TI2 PZ STG |  |  | TokaiIncubatorMCL:1:1:P
77025292 | STXG INCB F/WLP/DSH F/PRIORSP400 100%CO2 |  |  | TokaiIncubatorQueensgate:1:1:P | TokaiQueensgateRiser:0:1:H
77025293 | STXF INCB F/WLP/DSH F/PRIOR SP400 5%CO2 |  |  | TokaiIncubatorQueensgate:1:1:P | TokaiQueensgateRiser:0:1:H
77025334 | CHAMBER RISER FOR SP400X |  |  | TokaiQueensgateRiser:1:1:P
77025323 | STG ADPTR F/WSKMX CHMBR F/TI2 MAN T-S-SR |  |  | StageUInsert:1:1:P | TokaiIncubator:1:1:H
77025353 | DRK BX INCU 100%CO2 F/NANO-Z100/200-N/N2 |  |  | TokaiIncubatorMCL:1:1:P
77025354 | DARK BX INCU 5%CO2 F/NANO-Z100/200-N/N2 |  |  | TokaiIncubatorMCL:1:1:P
77025358 | DARK BX INCU 5%CO2 F/NANO-Z500-N2 PIEZO |  |  | TokaiIncubatorMCL500:1:1:P
77025357 | DRK BX INCU 100%CO2 F/NANO-Z500-N2 PIEZ |  |  | TokaiIncubatorMCL500:1:1:P
77025355 | DARK BX INCU 100%CO2 F/PRIOR QGSPPIEZO |  |  | TokaiIncubatorQueensgate:1:1:P | TokaiQueensgateRiser:0:1:H
77025356 | DARK BX INCU 5%CO2 F/PRIOR QGSP PIEZO |  |  | TokaiIncubatorQueensgate:1:1:P | TokaiQueensgateRiser:0:1:H
77025383 | DARK BX INCU HYPOXIA/PRIOR QGSPPIEZO |  |  | TokaiIncubatorQueensgate:1:1:P | TokaiQueensgateRiser:0:1:H
77025361 | DARK BX W/ TEMP CONTROLLER |  |  | StageEInsert:1:1:P
77025289 | STXF INCB F/WLP/DSH TI/TI2 PZ;STG/XY MTZ |  |  | TokaiIncubatorMCL:1:1:P
Objective Piezo Z Focus drives
77046507 | FN1 PIEZO FCSNG ELMT W/M32&M25 ADPTRS | NIS-PIEZO? | MCL 450um objective piezo for Ni or FN1 focusing nosepiece. Replaces the nosepiece, single objective holder includes 25mm and 32mm adapters for CFI60 or CFI75, w/slot for DIC slider.
77038015 | PIFOC PIEZO NANO FOCUSING Z-DRV 2MM TRVL | NIS-PIEZO? | PI Nano 2000um objective piezo for Ni or FN1 focusing nosepiece (MBP70110) with USB controller. Does not support NIDAQ. 25mm and 32mm adapters for CFI60 and CFI75 objectives.
77038029 | OBJ PIEZO Z DRV 2MM TRV W/DRCT ANL INPUT | NIS-PIEZO? | PI Nano 2000um objective piezo for Ni or FN1 focusing nosepiece (MBP70110) with direct analog controller for NIDAQ support. 25mm and 32mm adapters for CFI60 and CFI75 objectives.
Emission Wheels (non-triggered)
77016098 | LAMBDA 10-B F/1 FLTR WHL & SS OR 2 SS |  | Can control 1 filter wheel and 1 shutter OR 2 shutters
77016088 | LAMBDA 10-3 MICROPROCESSOR, MULTI-FILTER
77016405 | 32 MM FW 6 POS F/TI2
77016341 | NIKON TI2 STAGE UP, BRACKETS ONLY |  | Use in combination with MED53210 (Ti2E stage up) or MED53220 (Ti2A/U stage up)
77016334 | FN1 EMISSION FW ADAPTER
77012425 | CNTRLR MAC6000 DC FW NEW INTERFACE
77012039 | MAC6000 DC FW 10 POS 32/25MM, NO SHUTTER
77012426 | CUSTOM EMISSION FLANGE F/NIE QUAD HEAD |  | Use this to attach to the NiE quad head front port (with above two parts)
Transmitted LEDs (broadband)-BR91 and BR92 stock, LUMENCOR DISCONTINUED
77060084 | PEKA BRD WHT SPECT LT SOURCE, TI CNTRL
MXA22163 | TI2-100WRC 100W REMOTE CABLE |  | For Ti/Ti2 control of PEKA 77060084.
RGB Transmitted LEDs-ONLY BR92 MODELS EXIST NOW, LUMENCOR DISCONTINUED
77060081 | LIDA LIGHT ENGINE, RGB CHANLS |  | RGB Illumination driver only compatible with pco.edge, Flash, Zyla, or DS-Qi2 cameras
Broadband EPI LEDs
77060137 | SOLA V-NIR LIGHT ENGINE GENIII 380-760NM | NIS-EXFO? | Six-source 380-760nm broadband epifluorescence source. Liquid light guide is INCLUDED, | LumencorLLGSource:1:1:P |  | AddPart:77060021
77060129 | SOLA SE V-NIR | NIS-EXFO? | SOLA SE V-nIR, 380760 nm unfiltered light output, light on/off andgraduated intensity control via control pod or computer. Includes DCpower supply, power cord and USB control cable. Note: 3mm LLG(77060021) must be purchased separately. | LumencorLLGSource:1:1:P
77060136 | SOLA U-NIR LIGHT ENGINE GENIII 350-760NM | NIS-EXFO? | Six-source 350-760nm broadband epifluorescence source. Liquid light guide is INCLUDED, | LumencorLLGSource:1:1:P |  | AddPart:77060021
77060135 | SOLA LIGHT ENGINE GEN III 350-680NM | NIS-EXFO? | Four-source 350-680nm broadband epifluorescence source. Liquid light guide is INCLUDED | LumencorLLGSource:1:1:P |  | AddPart:77060021
77060113 | SOLA SE U-NIR W/CONTROL POD | NIS-EXFO? | SOLA SE U-nIR, 350#760 nm unfiltered light output, light on/off andgraduated intensity control via control pod or computer. Includes DCpower supply and power cord, control pod (83-10007), and USB controlcable. Note: 3mm LLG (77060021) must be purchased separately. | LumencorLLGSource:1:1:P
77060144 | SOLA FISH GENIII LIGHT ENGINE | NIS-EXFO? | SOLA FISH GENIII Light Engine, 350#680 nm unfiltered output spectrumoptimized for fluorescence in situ hybridization (FISH) applications.Light on/off and graduated intensity control via control pod accessoryor computer. Includes 3 mm diameter LLG, USB control cable, DC powersupply and power cord. Purchase Control Pod (77060083) separately. | LumencorLLGSource:1:1:P |  | AddPart:77060021
77060074 | SOLA SE II 365 LIGHT ENGINE | NIS-EXFO? | SOLA SE II 365 Light Engine (365nm replaces 395nm), Solid State White-Light Excitation SubSystem; ships with Adapter to LLG only. LLG(77060021) must be ordered separately. Controlled via touch screentablet or computer. If controlled by NIS-Elements, MQS41920 driverrequired. If triggered, MQS41920, MQS41930, MQS42780 is needed. | LumencorLLGSource:1:1:P
77060086 | SOLA SE II 365 LIGHT ENGINE W/CNTRL POD | NIS-EXFO? | SOLA SE II 365 Light Engine (365nm replaces 395nm), Solid State White-Light Excitation SubSystem; ships with Adapter to LLG and Control Podonly. LLG (77060021) must be ordered separately. | LumencorLLGSource:1:1:P
77060049 | SOLA SM II LE FOOT PEDAL AND CABLE
77060139 | LUMENCOR LLG 3MM 2M UV OPTIMIZED |  |  | LLG:1:1:P | LumencorLLGSource:1:1:H
77060021 | LUMENCOR LIQUID LIGHT GUIDE; 3MM, 2M VIS |  |  | LLG:1:1:P | LumencorLLGSource:1:1:H
77060101 | LUMENCOR LLG 3MM 2M REINFORCED JACKET |  | Metal jacketed LLG for extra protection. | LLG:1:1:P | LumencorLLGSource:1:1:H
77013515 | X-CITE XYLISII XT730S W/1.5M LLG W/ADPTR | NIS-EXFO? |  | XciteLLGSource:1:1:P |  | AddPart:77013360
77013510 | X-CITE XYLIS II, XT730S W/1.5M LLG | NIS-EXFO? |  | XciteLLGSource:1:1:P |  | AddPart:77013360
77013517 | X-CITE XYLIS II, XT730L W/1.5M LLG | NIS-EXFO? | X-Cite XYLIS II, XT730L with 1.5m LLG, without adapter. Includes: X-Cite XYLIS II (XT730L), SpeedDIAL remote, Accessory Kit (tools, USBcable, power cord), 3x1500mm LLG | XciteLLGSource:1:1:P |  | AddPart:77013360
77013498 | X-CITE 365NM XYLIS SYSTEM W/3MM 1.5M LLG | NIS-EXFO? | X-Cite XYLIS 365nm System. Includes XYLIS unit (XT720S), SpeedDIALremote, 3mmx1500mm Liquid Light Guide, Quick Start Guide, Hex Key, USBCable, Power Cord. | XciteLLGSource:1:1:P |  | AddPart:77013360
77013360 | X-CITE XLED1 - 1.5M LIQUID LIGHT GUIDE | NIS-EXFO? | Replacement LLG for Excelitas | LLG:1:1:P | XciteLLGSource:1:1:H
77013505 | X-CITE MINI PLUS XMPS 365 | NIS-EXFO?
77013506 | X-CITE MINI PLUS XMPL 385 | NIS-EXFO?
77013516 | NIKON FL-2 EPI FLANGE FOR X-CITE MINI+ |  | To fit HMX style (like 120 LED/Mini +) onto Nikon new style epi illuminator
Other Lumencor Light Sources
77060140 | CELESTA VBCTGRN W/DESPECKLING | LaserSafety,NIS-WLSWITCHER? | Lumencor CELESTA laser light engine - seven independently controlledlasers VBCTGRN with despeckling, ~800 mW per line. Integrated filteringincludes 405/12, 446/12, 476/12, 518/12, 545/12, 637/12, and 748/12.Includes 1.5 mm core optical SMA fiber with SMA and fiber-to-3mm LLGadapter, power supply and cord, and USB and ethernet control cables.Class 4 laser safety required. | CelestaLightSource:1:1:P | InterlockPhone:1:1:H
77060147 | SPCTRIII 348/380/440/475/510/555/575/635 | LaserSafety,NIS-WLSWITCHER? | SPECTRA III light engine: 348/30, 380/20, 440/20, 475/28, 510/25,555/28, 575/25, 635/22. Includes USB to RS232 cable,24VDC/9.2A powersupply, and power cord. LLG sold separately (77060139 for UV and77060021 for VIS). | LumencorLLGSource:1:1:P | InterlockPhone:1:1:H
77060120 | SPCTR III 390/440/475/510/555/575/635/74 | LaserSafety,NIS-WLSWITCHER? | SPECTRA III, VBCTGYRnIR Light Source with 390/22,440/20, 475/28, 510/25, 555/28, 575/25, 637/12 and 748/12 bandpassfilters. Ethernet, serial (USB B and RS232) and TTL (DB15) controlports. Includes USB control cable, DC power supply and cord. For usewith 3 mm diameter LLG (LLG sold separately). | LumencorLLGSource:1:1:P | InterlockPhone:1:1:H
77060138 | SPCTRIII 348/380/434/475/511/575/637/748 | LaserSafety,NIS-WLSWITCHER? | SPECTRA III UV,V,B,C,T,Y,R,nIR light engine: 348/30, 380/20, 434/21,475/28, 511/16, 575/25, 637/12, 748/12. Includes USB to RS232 cable,24VDC/9.2A power supply, and power cord. LLG sold separately (77060139for UV and 77060021 for VIS/NIR). | LumencorLLGSource:1:1:P | InterlockPhone:1:1:H
77060148 | AURA III 348/380/475/555/635 | NIS-WLSWITCHER? |  | LumencorLLGSource:1:1:P
77060124 | AURA III - UCGRNIR LIGHT ENGINE | NIS-WLSWITCHER? | Requires MQS41930 wavelength switcher and LLG. | LumencorLLGSource:1:1:P
77060130 | RETRA FURA III F/348/380NM | NIS-WLSWITCHER? | Gen III style light engine for FURA, 348/30 and 380/20 output. Includes 2 m UV optimized LLG. Direct replacement for original RETRA FURA. | LumencorLLGSource:1:1:P | AddPart:77060139
Intensilight
MBF72655 | C-HGFI INTENSILIGHT HG ILLUMINATOR |  |  | Intensilight:1:1:P | IntensilightLamp:1:1:H | PowerCord:1:1:H
MBF72665 | C-HGFIE INTENSILIGHT HG ILLUMINATOR MTRZ |  |  | Intensilight:1:1:P | IntensilightLamp:1:1:H | PowerCord:1:1:H | IntensilightController:1:1:H
MBF75600 | C-HGF IE-C HG CONTROLLER INTENSILIGHT |  |  | IntensilightController:1:1:P
MBF74650 | C-LHGF I HG INTENSILIGHT LAMP |  |  | IntensilightLamp:1:1:P
MBF71610 | C-HGF-IF15 HG FIBER 1500MM LLG |  |  | LLG:1:1:P | Intensilight:1:1:H
MBF71630 | C-HGF IF30 HG FIBER 3000MM LLG |  |  | LLG:1:1:P | Intensilight:1:1:H
NIDAQ Control
77023034 | NI-BB PLUS ANALG TRIGRD DEVCE HUB W/6323 | Triggering,NIS-TRIGGERING? | Includes NI-BB+, cable set, and NI PCIe6323 NIDAQ card. The NI‐BB+ is not compatible with | NIDAQKit:0:1:H | 6323Port:1:1:H | AddPart:77013248 | LUNPort:0:1:H | BNCBB:0:4:H
77013248 | PCIE-6323 NIDAQ CARD | Triggering,NIS-TRIGGERING? |  | PCIEport:1:1:P | 6323Port:1:1:P
77013233 | PCIE-6353 DAQ CARD 32AI AND 48 I/O LINES
77013240 | NIIPXI-1033 CONTROLLER
77013241 | NII PXI 6723 DAQ, LIMITED AVAILABILITY
77012523 | TRIG CBL/D-LEDI/HD15 LED OUTPUT F/NIBB+ |  | DLEDI trigger cable
77060082 | 4 CHANL BREAK OUT CABLE F/LIDA |  | LIDA trigger cable
77023027 | DEMO NI-BB TRIGGERED DEVICE HUB W/6323 | Triggering,NIS-TRIGGERING? | Includes NI PCIe6323 | NIDAQKit:1:1:H | 6323Port:1:1:H | AddPart:77013248
77023028 | DEMO NI-BB TRIGGERED DEVICE HUB | Triggering,NIS-TRIGGERING? | Does not include NIDAQ card: can be added to existing PXI6229 or PCIe 6353 | NIDAQKit:1:1:H | 6323Port:1:1:H
77010584 | NI DAQ HW BOX
77013235 | BNC CONNECT BLK F/NI6723 W/24 AO BNC CON
77013227 | NI SHC68-68-EPM SHIELDED CABLE
77013205 | CUSTOM CABLE 6723 DAQ TO 16 BNC
77013238 | NIDAQ ACCESSORY KIT |  |  | MISC:0:1:P | AddPart:77012801,77012801,77012801,77012801,77012801,77012801,77012801,77012801
77012801 | 6 FOOT MALE TO MALE BNC CABLE | BNC |  | BNC:0:1:P | BNCBB:0:1:P
uDarkRoom
77023047 | MICRO-DARK ROOM ENCLOSURE
Ti2 Hardware Development Kit
77013842 | TI2 HDK EPI PORT 30MM CAGE ADAPTER PLATE
77013843 | TI2 HDK EPI PORT 60MM CAGE ADAPTER PLATE
77013844 | TI2 HDK LAPP PRT ADPTR RNG F/30MM CAGE
77024092 | TI2HDK BASEPLATE STAGEUP
iLAS
77095005 | ILAS2 TIRF/PHOTOSTIM ILLUMIN FOR TI/TI2 | TIRF,Triggering,NIS-ILAS? | iLAS2 TIRF/Photostimulation illuminator for Ti/Ti2: two independentscanners in a compact illumination device compatible with LAPPilluminators. Requires an LAPP main branch. Includes controller, lasersafety cover, interlock connector box, PC interface connection, andcables. | LAPPSubBranchU¦LAPPSubBranchE:1:1:P | A1out¦A2out¦A3out¦A4out:1:1:H | A1out¦A2out¦A3out¦A4out:0:1:H | BNC:0:4:H | AddPart:77023046,MEE59177
97062 | BACK APERTURE CAM F/TI2 PHOTO BASE PORT | NIS-BACKAPCAM?,BackApertureCamera |  | CmountCamera:1:1:P | AddPart:77010612,MBB63435
Bits n Bobs
76998 | LENS TISSUE BOOK 50 SHEETS 4X6 |  |  | MISC:0:1:P
77002 | IMMERSION OIL TYPE A 1/4 OZ |  | Good for UV and Visible fluorescence. Not recommended for PFS or Super Resolution. | MISC:0:1:P
77004 | IMMERSION OIL TYPE B 1 OZ |  |  | MISC:0:1:P
77005 | IMMERSION OIL TYPE 37 1 OZ |  |  | MISC:0:1:P
77006 | IMMERSION OIL HF - 1 OZ NON-FLUORESCING |  |  | MISC:0:1:P
77007 | IMMERSION OIL LDF 1 OZ NON-FLUORESCING |  |  | MISC:0:1:P
77010 | IMMERSION OIL TYPE A 1 OZ |  | Good for UV and Visible fluorescence. Not recommended for PFS or Super Resolution. | MISC:0:1:P
79035 | POWER CORD (LEAD FREE UL LISTED) |  |  | PowerCord:0:1:P
79036 | POWER CORD R ANGLE (LEAD FREE UL LISTED) |  |  | PowerCord:0:1:P
91002 | INSTALLATION KIT |  | newer style accessory kit | MISC:0:1:P | AddPart:77155115
97050 | USB 2.0 A-B DEVICE CABLE 15' |  | Required for Elements Control Ti2-E or A | USBCord:0:1:P
77012942 | TRIPP LITE LC2400 LINE CONDITIONER |  |  | MISC:0:1:P
77013245 | USB3.0 A TO MICROB ANGLD W/LCK SCRW, 3M |  | Can be used to connect to USB 3 ss port (blue) on any HP workstation that has one available, or with USB 3 type A board | USB30AtoMicroB:0:1:P
77013253 | USB 3.2 GEN2X1 CABLE 10GBPS TRANSF USB-C |  | USB-C cable for Digital Sight 10 | USB32toUSB32:0:1:P
77013254 | USB 3.1 C TO MICRO B ANGLD CBLE F/DS-50M |  | Can be used to direct connect with HP workstations that have a C-port | USB31CtoMicroB:0:1:P
77013809 | SMALL CLAMPING FORK, 1.25" | ClampingFork |  | AXClamp:0:1:P
77013868 | CLAMPING FORK 1/4"-20 F/IMPERIAL |  |  | MISC:0:1:P
77013869 | CLAMPING FORK M6 X 1.0 F/METRIC |  |  | MISC:0:1:P
77019692 | PC ACCESSORY KIT |  | older style accessory kit | MISC:0:1:P | AddPart:77155115
77049093 | P15 TC AIR COMPRESSOR FOR VIT, 110V |  | Standard quiet compressor with oil. | MISC:0:1:P
77049300 | PC120/4 C QUIET OIL FREE COMPRESSOR |  |  | MISC:0:1:P
77050311 | 750x900MM; 30X36 VIT PACKAGE |  |  | MISC:0:1:P
77050312 | 900x900MM; 36X36 VIT PACKAGE |  |  | MISC:0:1:P
77050313 | 900x1200MM; 36X48 VIT PACKAGE |  |  | MISC:0:1:P
77050314 | 900x1500MM; 36X60 VIT PACKAGE |  |  | MISC:0:1:P
77150311 | 750x900MM; 30X36 VIT PACKAGE |  |  | MISC:0:1:P
77150312 | 900x900MM; 36X36 VIT PACKAGE |  |  | MISC:0:1:P
77150313 | 900x1200MM; 36X48 VIT PACKAGE |  |  | MISC:0:1:P
77150314 | 900x1500MM; 36X60 VIT PACKAGE |  |  | MISC:0:1:P
77150319 | SMOOTHTOP 30X48 VIT PACKAGE |  |  | MISC:0:1:P
77150320 | SMOOTHTOP 30X36 VIT PACKAGE |  |  | MISC:0:1:P
77049170 | VIB ISO TBL 30X48 CLEAN BENCH SMOOTH TOP |  |  | MISC:0:1:P
77049609 | VIB ISO PLATFORM TI-U, TI-U/B, TI-E,TI-E |  |  | MISC:0:1:P
77049634 | ISOLATION PLATFORM 16X20IN LAMINATE |  |  | MISC:0:1:P
77049814 | OPTICAL TABLE 6'X6"X8" F/AXMP |  |  | MISC:0:1:P
77049815 | VIB ISO PLATFORM 12X18 IN 10-60 LBS |  |  | MISC:0:1:P
77050322 | 750x750MM; 30X30 VIT PACKAGE |  |  | MISC:0:1:P
77050318 | Vibrations Isolation Faraday Table Package 900x1200mm; 36x48 with M6 on25mm tapped top, 59mm or 2"" thick. |  |  | MISC:0:1:P
77050315 | 900X1800MM; 36X72 VIT PACKAGE |  | Recommended for SoRa | MISC:0:1:P
77060083 | LE CNTRLR POD F/SPECTX/SOLA SE/MIRA/LIDA |  | Optional pod for manual control.
77060116 | 9-CHANNEL BNC BREAKOUT CABLE, 2M
JBAM-MISC | MISC |  |  | MISC:0:99:H
MXA22149 | S-TI2-EXT EXTERNAL TRIGGER CABLE |  | Ti2 trigger cable
MXA22168 | 30CC NON-FLUORESCING IMMERSION OIL F |  |  | MISC:0:1:P
MXA22179 | 30CC SILICONE IMMERSION OIL |  |  | MISC:0:1:P
MXA22202 | IMMERSION OIL N US 8CC |  |  | MISC:0:1:P
MXA22203 | IMMERSION OIL N 50CC |  |  | MISC:0:1:P
MXA22204 | IMMERSION OIL N US 480CC |  |  | MISC:0:1:P
92084 | ECLIPSE LARGE NYLON COVER 14X26X32 |  |  | MISC:0:1:P
77012940 | 4PRT USB PCIE 10GBPS 2XUSBC 2XUSBA 2CNTR |  | Card for DS-10 | USB4PortCard:1:1:P
MXA22001 | CFI60 OBJECTIVE LENS CASE |  |  | MISC:0:1:P
77013066 | LENS CLEANING SOLUTION 1 OZ SPRAY BOTTLE |  |  | MISC:0:1:P
77013067 | LENS CLEANING SOLUTION 6 OZ BOTTLE |  |  | MISC:0:1:P
91155 | NULL MODEM CABLE 10 FT |  |  | MISC:0:1:P
MQF52055 | AC ADAPTER L3 U3-NC |  |  | ACAdapter:1:1:P | PowerCord:1:1:H
MQF52056 | AC ADAPTER L3 U3-NC |  |  | ACAdapter:1:1:P | PowerCord:1:1:H
MQF52057 | AC ADAPTER |  |  | ACAdapter:1:1:P | PowerCord:1:1:H

## Cameras
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
Third Party Monochrome-PCIe Card sCMOS
77054128 | ORCA-QUEST2-KIT |  |  | CmountCamera:1:1:P | OrcaQuestInterface:0:1:H
77054121 | Hamamatsu ORCA Quest qCMOS camera. C-mount with USB 3.0 card andcable,power supply, and trigger cable. |  |  | CmountCamera:1:1:P | OrcaQuestInterface:0:1:H
77054122 | COAXPRESS CARD/CABLE KIT F/QUEST |  |  | OrcaQuestInterface:1:1:P | PCIEport:1:1:P
77054124 | ORCA FIRE QCMOS CAMERA C-MOUNT |  |  | CmountCamera:1:1:P | OrcaFireInterface:0:1:H
77054125 | COAXPRESS CARD/CABLE KIT F/ORCA FIRE |  |  | OrcaFireInterface:1:1:P | PCIEport:1:1:P
77018292 | PRIME 95B 25MM CMOS CAMERA |  | F-Mount, 2 year warranty, 25mm FOV | FmountCamera:1:1:P
77018291 | PRIME 95B CMOS 1200X1200 41FPS 16BIT |  | C-Mount; 2 year warranty, 18.7 mm FOV | CmountCamera:1:1:P
77054115 | ORCA-FUSION GEN-III SCMOS CAMERA C-MT |  | USB 30 FPS. Requires NIS-Elements 5.20.01 or greater. Add 77054116 to enable 89 FPS. | CmountCamera:1:1:P | OrcaFusionInterface:0:1:H | OrcaFusionDual:0:1:H
77054116 | INTERFACE KIT F/ORCA FUSION |  | CoaXPress interface for Fusion, enables 89 FPS 2304 x 2304. | OrcaFusionInterface:1:1:P | PCIEport:1:1:P
77054123 | DUAL ORCA FUSION/FUSION BT INTERFACE KIT | NIS-DUALCAM? |  | OrcaFusionDual:2:2:P | PCIEport:1:1:P
77054101 | ORCA-FLASH 4.0 SCI CMOS V3; USB 3.0 |  | USB 30 FPS. Add 77054102 to enable 100 FPS. | CmountCamera:1:1:P | OrcaFlashInterface:0:1:H
77054102 | CAMERA LINK OPTION FOR 77054101, 100FPS |  | Camera Link upgrade kit to 100fps for Flash 4.0 V3 | OrcaFlashInterface:1:1:P | PCIEport:1:1:P
77018299 | PRIME BSI SCMOS 2048X2048 C-MT |  | C-Mount; 2 year warranty. Back illuminated sCMOS. | CmountCamera:1:1:P
77018301 | IRIS 15 SCMOS CAM 5056X2968 32FPS@16BITS |  | F-Mount, 2 year warranty, 25mm FOV. | FmountCamera:1:1:P
Third Party Monochrome-USB sCMOS
77054118 | ORCA-FUSIONBT SCMOS USB3.0 CARD C-MT |  |  | CmountCamera:1:1:P | OrcaFusionInterface:0:1:H | OrcaFusionDual:0:1:H
77018306 | KINETIX 10MP CMOS CAMERA 3200X3200 |  | This is F-mount only, ~30mm FOV. | FmountCamera¦TmountCamera:1:1:P
77026135 | SONA 4.2B11 SCMOS 2048x2048 4MP |  |  | FmountCamera:1:1:P
77018310 | KINETIX 22 BACK-ILLUM SCMOS CAM C-MT |  | This is C-mount only, 22mm FOV. | CmountCamera:1:1:P
77026136 | SONA 2.0B-11 sCMOS 1400x1400 2MP |  |  | CmountCamera:1:1:P
77018305 | PRIME BSI EXPRESS CMOS 2048x2048 |  |  | CmountCamera:1:1:P
77054126 | ORCA-FLASH 4.0 LT3 SCMOS; USB 3.0 40FPS |  |  | CmountCamera:1:1:P
77067016 | PCO.EDGE 4.2BI USB3 SCMOS CAMERA |  |  | CmountCamera:1:1:P
77067008 | PCO.EDGE 4.2Q HIGH QE USB SCMOS CAMERA |  |  | CmountCamera:1:1:P
77067009 | PCO.EDGE 4.2LT USB3 SCMOS CAMERA |  |  | CmountCamera:1:1:P
77067012 | PCO.PANDA USB3.1 SCMOS CAMERA |  |  | CmountCamera:1:1:P
77018309 | MOMENT CMOS CAM 3200x2200 50FP@12BIT |  |  | CmountCamera:1:1:P
Third Party Monochrome-EMCCDs
77092001 | PROEM-HS 512BX3, 512X512, 16 UM PIXEL |  |  | CmountCamera:1:1:P
77092002 | PROEM-HS 1024BX3, 1024X1024, 13 UM PIXEL |  |  | CmountCamera:1:1:P
77018286 | EVOLVE DELTA 512X512 EMCCD MONO, AC |  |  | CmountCamera:1:1:P
77026122 | IXON LIFE 888 EMCCD 1024X1024 13UM PIXEL |  |  | CmountCamera:1:1:P
77026047 | IXON ULTRA DU897 EMCCD |  |  | CmountCamera:1:1:P
77026121 | IXON LIFE 897 EMCCD 512X512 16 UM PIXEL |  |  | CmountCamera:1:1:P
Nikon DS-Fi3 Color Camera
MQA18000 | DS-FI3 MICROSCOPE CAMERA |  |  | CmountCamera:1:1:P | ACAdapter:1:1:H | USB30AtoMicroB:0:1:H | Tablet:0:1:H
97060 | DS-FI3 CAM PKG W/USB-C/TABLET/KYBD |  | DS-Fi3 package with tablet, keyboard cover, pen, USB cable, power supply, and NIS-Elements L software license | AddPart:MQA18000,MQF52055
97061 | NIS-L TABLET SUPPORT KIT |  | Includes Surface Bluetooth mouse,Surface dock with 2 mini-DisplayPorts, 4 USB3.0 ports, 1 audio out, and1 gigabit ethernet port. Also includes mini-DisplayPort to HDMI adapter,15' HDMI cable, 6' mini-DisplayPort to DisplayPort cable, and 25' ethernet cable | Tablet:1:1:P
Nikon DS FX format Color CMOS
MQA22000 | DIGITAL SIGHT 50M CAMERA |  |  | FmountCamera:1:1:P | ACAdapter:1:1:H | USB30AtoMicroB¦USB31CtoMicroB:0:1:H
MQA20000 | DIGITAL SIGHT 10 CAMERA |  | Use recommended camera adapter MQD43050 DS-F1.8 F-MOUNT ADAPTER 1.8X | FmountCamera:1:1:P | ACAdapter:1:1:H | USB4PortCard:0:1:H | USB32toUSB32:1:1:H
MXA22148 | S-TI2-CTC CAMERA TRIGGER CABLE |  | Trigger cable for Qi2/Ri2. | Qi2TriggerCable:1:1:P
MQA17500 | DS-QI2 MONOCHROME CMOS CAMERA |  |  | FmountCamera:1:1:P | ACAdapter:1:1:H | USB30AtoMicroB:1:1:H | Qi2TriggerCable:0:1:H
Nikon Digital Sight 1000
MQA19000 | DIGITAL SIGHT 1000 MICROSCOPE CAMERA |  | For standalone operation via HDMI or NIS-Elements F via USB. NOT compatible with NIS-Elements AR/BR/D. 0.55x relay lens recommended. | CmountCamera:1:1:P | ACAdapter:1:1:H
97063 | DS1000 CAM W/TBLT CMPTR/NIS-ELEMENTS L |  | Digital Sight 1000 package with tablet, keyboard cover, pen, USB cable, power supply, extension dock, and NIS-Elements L software license. Must add power cord. | CmountCamera:1:1:P
97064 | TBLT CMPTR/NIS-ELEMENTSL F/DS-FI3/DS1000
Prime Life Science Wifi Color Cameras
77021041 | 12MP IN-LINE CLR CAM W/SCREEN EiSiCiNi |  | Intervision N12 Plus S for Nikon Ei/Si: 12MP All in one, in-line withscreen | CmountCamera:1:1:P
77021042 | 12MP+ INLINE CLR CAM W/O SCREEN EiSiCiNi |  | Intervision N12 Plus for Nikon Ei/Si: 12 MP All in one, in-line, noscreen | CmountCamera:1:1:P
77021043 | 12MP IN-LINE CLR CAM W/SCREEN E100/E200 |  | Intervision N12 S for Nikon E100/E200: 12 MP All in one, in-line withscreen | CmountCamera:1:1:P
77021044 | 12MP IN-LINE CLR CAM W/O SCREEN EiSiCiNi |  | Intervision N12 for Nikon Ei/Si: 12 MP All in one, in-line, no screen | CmountCamera:1:1:P
77021045 | 12MP IN-LINE CLR CAM SMZ800/1270 |  |  | CmountCamera:1:1:P
77021046 | 12MP IN-LINE CLR CAM trinocular model with screen |  |  | CmountCamera:1:1:P
77021047 | 12MP IN-LINE CLR CAM for Ts2 W/O SCREEN |  |  | CmountCamera:1:1:P
77021048 | 12MP IN-LINE CLR CAM W/ SCREEN |  |  | CmountCamera:1:1:P
77021049 | 12MP IN-LINE MONO CAM W/ SCREEN |  |  | CmountCamera:1:1:P
Image Splitters
77054107 | W-VIEW GEMINI IMAGE SPLITTER F/2 CAMERAS | NIS-DUALCAM? |  | SidePortAdapter:1:1:P | CmountCamera:2:2:H
77072064 | TI2 DIRCT SIDEPRT MT F/CAIRN SPLIT RANGE |  |  | SidePortAdapter:1:1:P | TwinCam:1:1:H
77072081 | TWINCAM DUAL CAM SPLITTER W/RECT APERTUR | NIS-DUALCAM? |  | TwinCam:1:1:P | TwinCamCmount:2:2:P
77072084 | C-MT OUTPUT CLAMP F/CAIRN SPLITTER RANGE |  |  | TwinCamCmount:1:1:H | CmountCamera:1:1:H
Couplers
77026139 | F MT TO C MT CNVRTR F/SONA 4.2BB11 SCMOS
MED53130 | T-BPA PHOTO ADAPTER |  | Installs in microscope sideport | CmountAdapter:0:1:P | BPAtoCmountAdapter:0:1:H
MED53135 | TI-BDTV D-SLR TV TUBE F/ F-MOUNT ADAPTER |  | Installs in place of microscope sideport
MED53140 | TI-BDTV2 F-TV TUBE F-MOUNT |  | Required for MQD42056 0.55x on Ti2 side port. | SidePortAdapter:0:1:P | FmountAdapter:0:1:H
MQD120AA | C MOUNT CAMERA ADAPTER 1X RELAY |  | Phototube
MQD120AF | C-MOUNT TV ADAPTER NEW |  | Phototube
MQD42005 | C-DA C-MOUNT/ISO CAMERA ADAPTER 1X |  | Reduces FOV to 22mm | CmountAdapter:0:1:P | CmountCamera:0:1:H
MQD42045 | C MOUNT CAMERA ADAPTER 0.45X-NC |  | Phototube | BPAtoCmountAdapter:0:1:P | CmountCamera:0:1:H
MQD42066 | C MOUNT CAMERA ADAPTER 0.6X-NC |  | Phototube | BPAtoCmountAdapter:0:1:P | CmountCamera:0:1:H
MQD42055 | C-0.55 DS RELAY LENS 0.55X |  |  | FmountAdapter:0:1:P | CmountCamera:0:1:H
MQD42056 | C-0.55 DS RELAY LENS 0.55X NC |  |  | FmountAdapter:0:1:P | CmountCamera:0:1:H
MQD42075 | C-0.7X DS RELAY LENS 0.70X NC |  | ISO port; Most Common | CmountAdapter:0:1:P | CmountCamera:0:1:H
MQD42120 | VM C MOUNT TV RELAY LENS 2.5X |  | Phototube | CmountAdapter:0:1:P | CmountAdapter:0:1:H
MQD42121 | VM C MOUNT RELAY LENS 2.5X SIM |  | ISO port | CmountAdapter:0:1:P | CmountAdapter:0:1:H
MQD42140 | VM C MOUNT TV RELAY LENS 4X |  | ISO port | CmountAdapter:0:1:P | CmountAdapter:0:1:H
MQD43000 | DS-F F-MOUNT ADAPTER 1X |  |  | FmountAdapter:0:1:P | FmountCamera¦F18Adapter:0:1:H
MQD43020 | DS-F2.5 F-MOUNT ADAPTER 2.5X |  |  | FmountAdapter:0:1:P | FmountCamera:0:1:H
MQD43050 | DS-F1.8 F-MOUNT ADAPTER 1.8X |  | For use wth DS-10, DS-50M | F18Adapter:0:1:P | FmountCamera:0:1:H
MQD44000 | DS10-IRCF MOTORIZED ADAPTER |  | Optional motorized adapter for switching "color" and "mono/IR" mode on DS-10. | FmountAdapter:0:1:P | F18Adapter:0:1:H
MQD51045 | V-ATV C-MOUNT ADAPTER 0.35X-NC |  | Phototube
MQD52010 | TI2-P-CCA CENTERING C-MOUNT ADPTR W/TOOL
MBB63435 | LV-TV-TV ADAPTER-CH |  |  | FmountAdapter:0:1:P | CmountAdapter:0:1:H
MBB63430 | LV-TV-TV ADAPTER |  |  | Alias:MBB63435
MXA29015 | D-SLR SERIES F MOUNT ADAPTER-NC |  | Not Compatible with Qi2/Ri2
JBAM-Ti2SidePort | Side Port Adapter for Ti2 |  |  | SidePortAdapter:0:1:P | CmountAdapter:0:1:H

## LUNF
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
77097004 | LUNF 405NM 50MW 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097005 | LUNF 405NM 100MW 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097006 | LUNF 488NM 100MW 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097007 | LUNF 488/561NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097008 | LUNF 440/514NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097009 | LUNF 405/488/561NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097010 | LUNF 405/488/561/640NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097011 | LUNF 405/488NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097012 | LUNF 532/561/640NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097013 | LUNF 488/561NM 1F AND 405NM 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097014 | LUNF 488NM 1F AND 561NM 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097015 | LUNF 405/488/561NM 1F AND 405NM 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097016 | LUNF 561NM 1F AND 640NM 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097017 | LUNF 488NM 1F 561NM 2F 640NM 3F | LaserSafety,Triggering |  | Alias:LUNF3Fiber
77097018 | LUNF 440NM 1F 514NM 2F 594NM 3F | LaserSafety,Triggering |  | Alias:LUNF3Fiber
77097019 | LUNF 405NM 1F 488NM 2F 561NM 3F | LaserSafety,Triggering |  | Alias:LUNF3Fiber
77097020 | LUNF 405NM 1F 488NM 2F 561NM 3F 640NM 4F | LaserSafety,Triggering |  | Alias:LUNF4Fiber
77097024 | LUNF 488/561/640 NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097025 | LUNF 488/561/640 NM 1F 405 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097026 | LUNF 445/488/561/640NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097027 | LUNF 445/488/561NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097028 | LUNF 445/514/561/640NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097029 | LUNF 445/514/561NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097030 | LUNF 532/640NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097031 | LUNF 405/445/488/561 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097032 | LUNF 445/488/561 NM 1F 405 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097033 | LUNF 405/445/561 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097034 | LUNF 405/514NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097035 | LUNF 515NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097036 | LUNF 561/640NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097037 | LUNF 640NM 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097039 | LUNF 405/488/514/561NM 1F FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097040 | LUNF 405/514/561NM 1F FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097041 | LUNF 488/514/561NM 1F FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097042 | LUNF 488/514/561/640NM 1F FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097043 | LUNF 488/647 NM 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097044 | LUNF 488/561/640 NM 1F APC 405 2F APC | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097045 | LUNF 488 1F FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097047 | LUNF 405 1F 473 2F | LaserSafety,Triggering |  | Alias:LUNF2Fiber
77097048 | LUNF LASER COMB 405/488/561/640 CSUX/W-1 | LaserSafety,Triggering |  | P1¦P2¦P3¦P4:1:1:H | AddPart:JBAM-PC8Laser1 | InterlockLUNF:1:1:H | LUNPort:0:1:P | AddPart:77012509,JBAM-LUNFGenderChanger
77097049 | LUNF LASER COMB 445/488/514/561 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097050 | LUNF LASER COMB 405/445/488/532 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097051 | LUNF STORM 405/488/561/640 1F APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097052 | LUNF LASER COMB 405/445/488 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097053 | LUNF 405NM 1F 473NM 2F 594NM 3F | LaserSafety,Triggering |  | Alias:LUNF3Fiber
77097057 | LUNF LASER COMB 445/514/594 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097059 | LUNF 405NM 50MW 1F UPGRADABLE | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097060 | LUNF LASER COMB 405/445 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097061 | LUNF LASER COMB 473 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097063 | LUNF LASER COMB 488/532 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097064 | LUNF 405/445/561/640 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097065 | LUNF LASER LAUNCH 561 1F | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097066 | LUNF LASR LAUNCH COMB 473 1F UPGRADEABLE | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097067 | LUNF 405/473 1F FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097068 | LUNF 405/473 1F FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097069 | LUNF 488/532/640 APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097070 | LUNF LASER COMB 405/488/561/640 20MW APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097073 | LUNF 488/532/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77097089 | LUNF HP 405/488/561/640 PC8 | LaserSafety,Triggering |  | Alias:LUNF1Fiber
77098010 | LUNF XL 405/488/561/640 PC8 405/50MW APC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098011 | LUNF XL 405/445/488/514/561/640 FC-PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098012 | LUNF XL 405/445/488/514/561/640PC8 405/50MWAPC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098013 | LUNF XL 405/445/488/514/561/594/640 FC-PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098014 | LUNF XL 405/488/514/532/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098015 | LUNF XL 405/488/514/561/594/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098016 | LUNF XL 405/488/514/561/640 473 2F FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098017 | LUNF XL 405/445/488/514/561/640,405/488 2F | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098018 | LUNF XL 405/488/445/514/561/594/640 405/488 | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098019 | LUNF XL 405/445/532/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098020 | LUNF XL 405/488/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098021 | LUNF XL 405/445/488/514/561/594/640, 4052F | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098022 | LUNF XL 405/488/514/532/561/594/640FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098023 | LUNF XL 405/445/488/532/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098024 | LUNF XL 405/445/488/532/561 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098025 | LUNF XL 405/445/488/514/561/640, 4052F | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098026 | LUNF XL 405/488/514/561/640 FC-PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098027 | LUNF XL 405/488/561/640 2FB TIRF/CSU | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098028 | LUNF XL 488/561 1 FBR FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098029 | LUNF XL 405/488/561/594/640 FC-PC8 FB | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098030 | LUNF XL 405/488/561/640 FC-PC8 | LaserSafety,Triggering
77098031 | LUNF XL 405/488/561/640 PC8; 405/473 APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098032 | LUNF XL 405/488/561/640 FC-PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098033 | LUNF XL 405/445/488/514/561/594/640 PC8, 4052F | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098034 | LUNF XL 405/488/561/640PC8 488/561&405 APC 3F | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098035 | LUNF XL 405/488/532/594/640 FC-PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098036 | LUNF XL 405/445/488/514/561/594/640 405/473/561 | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098038 | LUNF XL 405/445/488/514/561 PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098039 | LUNF XL 405/445/488/514/561/594/640 405/473F2 | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098040 | LUNF XL 405/445/488/514/561/640 405/473/561F2 | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098041 | LUNF XL 405/488/561/640 FC-APC 405F2 APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098042 | LUNF XL 405/445/488/514/561-PC8 405-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098043 | LUNF XL 405/488/561/640PC8-F1/APC-F2 405APC-F3 | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098044 | LUNF XL 405/488/561/594/640-PC8 405-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098045 | LUNF XL 405/445/488/561/594/640PC8405/488/561A | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098046 | LUNF XL 405/445/488/561/594/640 FC-PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
77098047 | LUNF XL 405/445/488/561/594/640PC8 488APC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098048 | LUNF XL 405/488/561/640-PC8 488-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098049 | LUNF XL 445/514/561/640,405/445/514/561 2F APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098050 | LUNF XL 405/488/561/640-PC8 405-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098051 | LUNF XL 405/488/561/640-PC8 405-APC UPG | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098052 | LUNF XL 405/488/532/561/594/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098053 | LUNF XL 405/488/561/640-PC8 405/488/561-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPCPC8
77098054 | LUNF XL 488/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098055 | LUNF XL 445/514/561/640, 405/445/514 APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098056 | LUNF XL 532/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098057 | LUNF XL 405/488/532/561/594/640 PC8 488/561/640 APC 405 APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098068 | LUNF XL 405/445/488/514/561/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098074 | LUNF XL 405/445/488/514/561/594/640 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098076 | LUNF XL 405/488/561/640 APC 405/50MW APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098079 | LUNF XL 405/488/561/640 laser combiner (100/100/100/125mW) FC-APC outputfiber. Includes USB cable, interlock | LaserSafety,Triggering |  | Alias:LUNFXLAPC
77098108 | LUNF XL 405 FC-APC 488 FC-APC | LaserSafety,Triggering |  | Alias:LUNFXLAPC2Fiber
77098110 | LUNF XL 445/488/514/561/594/640 PC8 | LaserSafety,Triggering |  | Alias:LUNFXLPC8
JBAM-APCLaser1 | APC Laser Line 1 |  |  | A1:1:1:P | A1out:0:1:P
JBAM-APCLaser2 | APC Laser Line 2 |  |  | A2:1:1:P | A2out:0:1:P
JBAM-APCLaser3 | APC Laser Line 3 |  |  | A3:1:1:P | A3out:0:1:P
JBAM-APCLaser4 | APC Laser Line 4 |  |  | A4:1:1:P | A4out:0:1:P
JBAM-PC8Laser1 | PC8 Laser Line 1 |  |  | P1:1:1:P | P1out:0:1:P
JBAM-PC8Laser2 | PC8 Laser Line 2 |  |  | P2:1:1:P | P2out:0:1:P
JBAM-PC8Laser3 | PC8 Laser Line 3 |  |  | P3:1:1:P | P3out:0:1:P
JBAM-PC8Laser4 | PC8 Laser Line 4 |  |  | P4:1:1:P | P4out:0:1:P
JBAM-LUNF1Fiber | LUNF1Fiber | LaserSafety,Triggering |  | A1¦A2¦A3¦A4:1:1:H | AddPart:JBAM-APCLaser1 |  |  | InterlockLUNF:1:1:H | LUNPort:0:1:P | AddPart:77012509,JBAM-LUNFGenderChanger
JBAM-LUNF2Fiber | LUNF2Fiber | LaserSafety,Triggering |  | A1¦A2¦A3¦A4:2:2:H | AddPart:JBAM-APCLaser1,JBAM-APCLaser2 |  |  | InterlockLUNF:1:1:H | LUNPort:0:1:P | AddPart:77012509,JBAM-LUNFGenderChanger
JBAM-LUNF3Fiber | LUNF3Fiber | LaserSafety,Triggering |  | A1¦A2¦A3¦A4:3:3:H | AddPart:JBAM-APCLaser1,JBAM-APCLaser2,JBAM-APCLaser3 |  |  | InterlockLUNF:1:1:H | LUNPort:0:1:P | AddPart:77012509,JBAM-LUNFGenderChanger
JBAM-LUNF4Fiber | LUNF4Fiber | LaserSafety,Triggering |  | A1¦A2¦A3¦A4:4:4:H | AddPart:JBAM-APCLaser1,JBAM-APCLaser2,JBAM-APCLaser3,JBAM-APCLaser4 |  |  | InterlockLUNF:1:1:H | LUNPort:0:1:P | AddPart:77012509,JBAM-LUNFGenderChanger
JBAM-LUNFXLPC8 | LUNFXLPC8 | LaserSafety,Triggering |  | P1¦P2¦P3¦P4:1:1:H | AddPart:JBAM-PC8Laser1 |  | BNCBB:0:4:H | InterlockLUNF:1:1:H | 6323Port:1:1:H | AddPart:77012509,JBAM-LUNFGenderChanger
JBAM-LUNFXLAPC | LUNFXLAPC | LaserSafety,Triggering |  | A1¦A2¦A3¦A4:1:1:H | AddPart:JBAM-APCLaser1 |  | BNCBB:0:4:H | InterlockLUNF:1:1:H | 6323Port:1:1:H | AddPart:77012509,JBAM-LUNFGenderChanger
JBAM-LUNFXLAPCPC8 | LUNFXLAPCPC8 | LaserSafety,Triggering |  | A1¦A2¦A3¦A4:1:1:H | P1¦P2¦P3¦P4:1:1:H | AddPart:JBAM-PC8Laser1,JBAM-APCLaser1 | BNCBB:0:4:H | InterlockLUNF:1:1:H | 6323Port:1:1:H | AddPart:77012509,JBAM-LUNFGenderChanger
JBAM-LUNFXLAPC2Fiber | LUNFXLAPC2Fiber | LaserSafety,Triggering |  | A1¦A2¦A3¦A4:2:2:H | AddPart:JBAM-APCLaser1,JBAM-APCLaser2 |  | BNCBB:0:4:H | InterlockLUNF:1:1:H | 6323Port:1:1:H | AddPart:77012509,JBAM-LUNFGenderChanger

## Fluorescence Filters
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
96121 | EF-4 AT DAPI/HOECHST/ALEXA FLUOR 350 |  | EF-4 AT DAPI/Hoechst/Alexa Fluor 350 Filter Set, Excitation: 375/28nm (361-389nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 415nm | FilterCube:0:1:P
96122 | EF-4 AT AQUA LP |  | EF-4 AT Aqua Long Pass Filter Set, Excitation: 420/40nm (400-440nm), Emission: 465nm Long Pass, Dichroic Mirror: 455nm | FilterCube:0:1:P
96123 | EF-4 AT ECFP/CERULEAN |  | EF-4 AT ECFP/Cerulean Filter Set, Excitation: 435/20nm (425-445nm), Emission: 480/30nm (465-495nm), Dichroic Mirror: 455nm | FilterCube:0:1:P
96124 | EF-4 AT FLUORO GOLD/CALCOFLUOR WHITE LP |  | EF-4 AT Fluoro Gold/Calcofluor White Long Pass Filter Set, Excitation: 375/28nm (361-389nm), Emission: 515nm Long Pass, Dichroic Mirror: 415nm | FilterCube:0:1:P
96125 | EF-4 AT GFP/FITC LP |  | EF-4 AT GFP/FITC Long Pass Filter Set, Excitation: 480/30nm (465-495nm), Emission: 515nm Long Pass, Dichroic Mirror: 505nm | FilterCube:0:1:P
96126 | EF-4 AT EGFP/FITC/CY2/ALEXA FLUOR 488 |  | EF-4 AT EGFP/FITC/Cy2/Alexa Fluor 488 Filter Set, Excitation: 480/30nm (465-495nm), Emission: 535/20nm (525-545nm). Dichroic Mirror: 505nm | FilterCube:0:1:P
96127 | EF-4 AT EYFP/VENUS/CITRINE |  | EF-4 AT EYFP/Venus/Citrine Filter Set, Excitation: 495/20nm (485-505nm), Emission: 540/30nm (525-555nm), Dichroic Mirror: 515nm | FilterCube:0:1:P
96128 | EF-4 AT TRITC/CY3 LP |  | EF-4 AT TRITC/Cy3 Long Pass Filter Set, Excitation: 540/25nm (527.5- 552.5nm), Emission: 575 Long Pass, Dichroic Mirror: 565nm | FilterCube:0:1:P
96129 | EF-4 AT TRITC/CY3/TagRFP/ALEXA FLUOR 546 |  | EF-4 AT TRITC/Cy3/TagRFP/Alexa Fluor 546 Filter Set, Excitation: 540/25nm (527.5-552.5nm), Emission: 605/55nm (577.5-632.5nm), Dichroic Mirror: 565nm | FilterCube:0:1:P
96130 | EF-4 AT TRITC (RED SHIFTED) |  | EF-4 AT TRITC (Red Shifted) Filter Set, Excitation: 540/25nm (527.5- 552.5nm), Emission: 620/60nm (590-650nm), Dichroic Mirror: 565nm | FilterCube:0:1:P
96131 | EF-4 AT TEXAS RED/mCHERRY/ALEXA FLUOR 59 |  | EF-4 AT Texas Red/mCherry/Alexa Fluor 594 Filter Set, Excitation: 560/40nm (540-580nm), Emission: 635/60nm (605-665), Dichroic Mirror: 600nm | FilterCube:0:1:P
96132 | EF-4 AT CY5/ALEXA FLUOR 647/DRAQ 5 |  | EF-4 AT CY5/Alexa Fluor 647/Draq 5 Filter Set, Excitation: 620/50nm (595-645nm), Emission: 690/50nm (665-715nm), Dichroic Mirror: 655nm | FilterCube:0:1:P
96134 | EF-4 DAPI HC SOLA HISN ZERO SHIFT |  | EF-4 DAPI HC SOLA HISN Zero Shift Filter Set, Excitation: 395/25nm (382.5-407.5nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 425nm | FilterCube:0:1:P
96147 | EF-4 HC FITC/TRITC |  | EF-4 HC FITC/TRITC Dual Band Filter Set | FilterCube:0:1:P
96148 | EF-4 HC FITC/TEXAS RED |  | EF-4 HC FITC/Texas Red Dual Band Filter Set | FilterCube:0:1:P
96149 | EF-4 HC DAPI/FITC/TRITC |  | EF-4 HC DAPI/FITC/TRITC Triple Band Filter Set | FilterCube:0:1:P
96150 | EF-4 HC DAPI/FITC/TEXAS RED |  | EF-4 HC DAPIFITC/Texas Red Triple Band Filter Set | FilterCube:0:1:P
96220 | C-FL AT DAPI/UV LP |  | CF-L AT DAPI/UV Long Pass Filter Set, Excitation: 375/28nm (361-389nm),Emission: 435nm Long Pass, Dichroic Mirror: 415nm | FilterCube:0:1:P
96221 | C-FL AT DAPI/HOECHST/ALEXA FLUOR 350 |  | CF-L AT DAPI/Hoechst/Alexa Fluor 350 Filter Set, Excitation: 375/28nm (361-389nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 415nm | FilterCube:0:1:P
96222 | C-FL AT AQUA LP |  | CF-L AT Aqua Long Pass Filter Set, Excitation: 420/40nm (400-440nm),Emission: 465nm Long Pass, Dichroic Mirror: 455nm | FilterCube:0:1:P
96223 | C-FL AT ECFP/CERULEAN |  | CF-L AT ECFP/Cerulean Filter Set, Excitation: 435/20nm (425-445nm), Emission: 480/30nm (465-495nm), Dichroic Mirror: 455nm | FilterCube:0:1:P
96224 | C-FL AT FLUORO GOLD/CALCOFLUOR WHITE LP |  | CF-L AT Fluoro Gold/Calcofluor White Long Pass Filter Set, Excitation:375/28nm (361-389nm), Emission: 515nm Long Pass, Dichroic Mirror: 415nm | FilterCube:0:1:P
96225 | C-FL AT GFP/FITC LP |  | CF-L AT GFP/FITC Long Pass Filter Set, Excitation: 480/30nm (465-495nm),Emission: 515nm Long Pass, Dichroic Mirror: 505nm | FilterCube:0:1:P
96226 | C-FL AT EGFP/FITC/CY2/ALEXA FLUOR 488 |  | CF-L AT EGFP/FITC/Cy2/Alexa Fluor 488 Filter Set, Excitation: 480/30nm (465-495nm), Emission: 535/40nm (515-555nm). Dichroic Mirror: 505nm | FilterCube:0:1:P
96227 | C-FL AT EYFP/VENUS/CITRINE |  | CF-L AT EYFP/Venus/Citrine Filter Set, Excitation: 495/20nm (485-505nm), Emission: 540/30nm (525-555nm), Dichroic Mirror: 515nm | FilterCube:0:1:P
96228 | C-FL AT TRITC/CY3 LP |  | CF-L AT TRITC/Cy3 Long Pass Filter Set, Excitation: 540/25nm (527.5-552.5nm), Emission: 575 Long Pass, Dichroic Mirror: 565nm | FilterCube:0:1:P
96229 | C-FL AT TRITC/CY3/TagRFP/ALEXA FLUOR 546 |  | CF-L AT TRITC/Cy3/TagRFP/Alexa Fluor 546 Filter Set, Excitation: 540/25nm (527.5-552.5nm), Emission: 605/55nm (577.5-632.5nm), Dichroic Mirror: 565nm | FilterCube:0:1:P
96230 | C-FL AT TRITC (RED SHIFTED) |  | CF-L AT TRITC (Red Shifted) Filter Set, Excitation: 540/25nm (527.5-552.5nm), Emission: 620/60nm (590-650nm), Dichroic Mirror: 565nm | FilterCube:0:1:P
96231 | C-FL AT TEXAS RED/mCHERY/ALEXA FLUOR 594 |  | CF-L AT Texas Red/mCherry/Alexa Fluor 594 Filter Set, Excitation: 560/40nm (540-580nm), Emission: 635/60nm (605-665), Dichroic Mirror: 600nm | FilterCube:0:1:P
96232 | C-FL AT CY5/ALEXA FLUOR 647/DRAQ 5 |  | CF-L AT CY5/Alexa Fluor 647/Draq 5 Filter Set, Excitation: 620/50nm (595-645nm), Emission: 690/50nm (665-715nm), Dichroic Mirror: 655nm | FilterCube:0:1:P
96240 | C-FL 385nm LED FILTER SET |  | C-FL 385nm LED-DAPI Filter Set, Excitation: 390/38nm (371-409nm), Emission: 475/90nm (430-520nm), Dichroic Mirror: 420nm | FilterCube:0:1:P
96241 | C-FL 455nm LED FILTER SET |  | C-FL 455nm LED-CFP Long Pass Filter Set, Excitation 448/23nm (436.5-459.5nm), Emission 472nm Long Pass, Dichroic Mirror: 465nm | FilterCube:0:1:P
96242 | C-FL 470nm LED FILTER SET |  | C-FL 470nm LED-FITC, GFP Filter Set, Excitation: 470/40nm (450-490nm), Emission: 534/55nm (506.5-561.5nm), Dichroic Mirror: 500nm | FilterCube:0:1:P
96243 | C-FL 505nm LED FILTER SET |  | C-FL 505nm LED-YFP Filter Set, Excitation: 496/29nm (481.5-510.5nm), Emission: 543/37nm (524.5-561.5nm), Dichroic Mirror: 520nm | FilterCube:0:1:P
96244 | C-FL 525nm LED FILTER SET |  | C-FL 525nm LED-Blue/Green Filter Set, Excitation: 525/50nm (500-550nm), Emission: 597/59nm (567.5-626.5nm), Dichroic Mirror: 560nm | FilterCube:0:1:P
96245 | C-FL 560nm LED FILTER SET |  | C-FL 560nm LED-TRITC, Cy3 Filter Set, Excitation: 550/50nm (527-575nm), Emission: 630/75nm (602.5-667.5nm), Dichroic Mirror: 600nm | FilterCube:0:1:P
96246 | C-FL 590nm LED FILTER SET |  | C-FL 590nm LED-mCherry, Texas Red Filter Set, Excitation: 561/75nm(523.5-598.5nm), Emission: 652/65nm (619.5-684.5nm), Dichroic Mirror: 610nm | FilterCube:0:1:P
96247 | C-FL 625nm LED FILTER SET |  | C-FL 625nm LED-Cy5 Filter Set, Excitation: 621/58nm (592-650nm), Emission: 706/73nm (669.5-742.5nm), Dichroic Mirror: 660nm | FilterCube:0:1:P
96300 | C-FL UV-1A |  | C-FL UV-1A | FilterCube:0:1:P
96302 | C-FL B-2A |  | C-FL B-2A, Excitation: 470/40nm (450-490nm), Emission: 515nm, Dichroic Mirror: 500nm. | FilterCube:0:1:P
96303 | C-FL BV/B CALCOFLUOR WHITE |  | C-FL BV/B Calcofluor White, Excitation: 425/40nm (405-445nm), Emission: 520nm, Dichroic Mirror: 475nm | FilterCube:0:1:P
96305 | C-FL G-1B |  | C-FL G-1B, Excitation: 546/10nm (541-551nm), Emission: 590nm Long Pass, Dichroic Mirror: 565nm | FilterCube:0:1:P
96306 | C-FL G-2A |  | C-FL G-2A, Excitation: 535/50nm (510-560nm), Emission: 590nm Long Pass, Dichroic Mirror: 565nm | FilterCube:0:1:P
96310 | C-FL UV-2E/C DAPI |  | C-FL UV-2E/C DAPI/Hoeschst (UV-3A), Excitation: 360/40nm (340-380nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 400nm | FilterCube:0:1:P
96311 | C-FL B-2E/C FITC |  | C-FL B-2E/C FITC/FLUO-3 (B-4A), Excitation: 480/30nm (465-495nm), Emission: 535/40nm (515-555nm), Dichroic Mirror: 505nm | FilterCube:0:1:P
96312 | C-FL G-2E/C TRITC |  | C-FL G-2E/C TRITC, Excitation: 540/25nm (528-553nm), Emission: 620/60nm (590-650nm), Dichroic Mirror: 565nm | FilterCube:0:1:P
96313 | C-FL Y-2E/C TEXAS RED |  | C-FL Y-2E/C Texas Red, Excitation: 560/40nm (540-580nm), Emission: 630/60nm (600-660nm), Dichroic Mirror: 595nm | FilterCube:0:1:P
96320 | C-FL FITC HYQ |  | C-FL FITC HYQ, Excitation: 480/40nm (460-500nm), Emission: 535/50nm (510-560nm), Dichroic Mirror: 505nm | FilterCube:0:1:P
96321 | C-FL TRITC HYQ |  | C-FL TRITC HYQ, Excitation: 545/30nm (530-560nm), Emission: 620/60nm (590-650nm), Dichroic Mirror: 570nm | FilterCube:0:1:P
96322 | C-FL TEXAS RED HYQ |  | C-FL Texas Red HYQ, Excitation: 560/55nm (532-587nm), Emission: 645/75nm (608-683nm), Dichroic Mirror: 595nm | FilterCube:0:1:P
96323 | C-FL CY3 HYQ |  | C-FL CY3 HYQ, Excitation: 545/30nm (530-560nm), Emission: 610/75nm (573-648), Dichroic Mirror: 570nm | FilterCube:0:1:P
96324 | C-FL CY5 HYQ |  | C-FL CY5 HYQ, Excitation: 620/60nm (590-650nm), Emission: 700/75nm (663 -738nm), Dichroic Mirror: 660nm | FilterCube:0:1:P
96341 | C-FL CYAN GFP |  | C-FL Cyan Fluorescent Protein (CFP), Excitation: 436/20nm (426-446nm), Emission: 480/40nm (460-500nm), Dichroic Mirror: 455nm | FilterCube:0:1:P
96342 | C-FL ENDOW GFP HYQ LP |  | C-FL Endow Green Fluorescent Protein (GFP) HYQ, 500 Long Pass Barrier, Excitation: 470/40nm (450-490nm), Emission: 500nm Long Pas s, Dichroic Mirror: 495nm | FilterCube:0:1:P
96343 | C-FL ENDOW GFP HYQ BP |  | C-FL Endow Green Fluorescent Protein (GFP) HYQ, 525 Bandpass Barrier, Excitation: 470/40nm (450-490nm), Emission: 525/50nm (500-5 50nm), Dichroic Mirror: 495nm | FilterCube:0:1:P
96345 | C-FL YELLOW GFP HYQ |  | C-FL Yellow Fluorescent Protein (YFP) BP HYQ, 535 Bandpass Barrier, Excitation: 500/20nm (490-510nm), Emission: 535/30nm (520-550 nm), Dichroic Mirror: 515nm | FilterCube:0:1:P
96350 | C-FL DAPI/FITC |  | C-FL DAPI/FITC (D-F), Excitation: 401-418nm,478-500nm, Emission: 450 -465nm, 512-560nm | FilterCube:0:1:P
96351 | C-FL FITC/TRITC |  | C-FL FITC/TRITC (F-R), Excitation: 475-494nm,545-565nm, Emission: 503nm-533nm, 582-622nm. | FilterCube:0:1:P
96352 | C-FL FITC/TEXAS RED |  | C-FL FITC/Texas Red (F-T), Excitation: 490-505nm, 560-585nm, Emission: 515-545nm, 600-652nm | FilterCube:0:1:P
96355 | C-FL DAPI/FITC/TRITC |  | C-FL DAPI/FITC/TRITC (D-F-R), Excitation: 385-400nm, 475-493nm, 545 -565nm, Emission: 450-465nm, 503-533nm, 582-622nm | FilterCube:0:1:P
96356 | C-FL DAPI/FITC/TEXAS RED |  | C-FL DAPI/FITC/Texas Red (D-F-T), Excitation: 395-410nm, 490-505nm, 560 -585nm, Emission: 450-475nm, 515-545nm, 600-652nm | FilterCube:0:1:P
96359 | C-FL DAPI HC SOLA HISN ZERO SHIFT |  | C-FL DAPI SOLA Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 395/25nm (382.5-407.5nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 425nm | FilterCube:0:1:P
96360 | C-FL DAPI HC HISN ZERO SHIFT |  | C-FL DAPI Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 350/50nm (325-375nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 400nm | FilterCube:0:1:P
96361 | C-FL CFP HC HISN ZERO SHIFT |  | C-FL CFP Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 436/20nm (426-446nm), Emission: 480/40nm (460-500nm), Dichroic Mirror: 455nm | FilterCube:0:1:P
96362 | C-FL GFP HC HISN ZERO SHIFT |  | C-FL GFP Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 470/40nm (450-490nm), Emission: 525/50nm (500-550nm), Dichroic Mirror: 495nm | FilterCube:0:1:P
96363 | C-FL YFP HC HISN ZERO SHIFT |  | C-FL YFP Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 500/20nm (490-510nm), Emission: 535/30nm (520-550nm), Dichroic Mirror: 515nm | FilterCube:0:1:P
96364 | C-FL DS RED HC HISN ZERO SHIFT |  | C-FL DS Red Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 545/30nm (530-560nm), Emission: 620/60nm (590-630nm), Dichroic Mirror: 570nm | FilterCube:0:1:P
96365 | C-FL TEXAS RED HC HISN ZERO SHIFT |  | C-FL Texas Red Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 560/40nm (540-580nm), Emission: 630/70nm (593-668nm), Dichroic Mirror: 585nm | FilterCube:0:1:P
96366 | C-FL CY5 HC HISN ZERO SHIFT |  | C-FL CY5 Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 620/60nm (590-650nm), Emission: 700/75nm (663-738nm), Dichroic Mirror: 660nm | FilterCube:0:1:P
96367 | C-FL CY7 HC HISN ZERO SHIFT |  | C-FL CY7 Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 710/75nm (672.5-747.5nm), Emission: 810/90nm (765-855nm), Dichroic Mirror: 760nm | FilterCube:0:1:P
96369 | C-FLL LFOV DAPI SOLA HISN SEM BL SET |  | C-FLL Large Field of View DAPI SOLA Filter Set, High-Signal-Noise, Semrock BrightLine®, Excitation: 378/52nm (352-404nm), Emission: 447/60nm (417-477nm), Dichroic Mirror: 409nm | FilterCubeLFOV:0:1:P
96370 | C-FLL LFOV DAPI HISN SEM BL FILTER |  | C-FLL Large Field of View DAPI Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 356/30nm (341-371nm), Emission: 447/60nm (417- 477nm), Dichroic Mirror: 409nm | FilterCubeLFOV:0:1:P
96371 | C-FLL LFOV CFP HISN SEM BL FIL |  | C-FLL Large Field of View CFP Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 438/24nm (426-450nm), Emission: 483/32nm (467- 499nm), Dichroic Mirror: 458nnm | FilterCubeLFOV:0:1:P
96372 | C-FLL LFOV GFP HISN SEM BL FILTER SET |  | C-FLL Large Field of View GFP/FITC/Cy2 Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 466/40nm (446-486nm), Emission: 525/50nm (500-550nm), Dichroic Mirror: 495nm | FilterCubeLFOV:0:1:P
96373 | C-FLL LFOV YFP HISN SEM BL FILTER SET |  | C-FLL Large Field of View YFP Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 500/24nm (488-515nm), Emission: 542/27nm (528.5-555.5nm), Dichroic Mirror: 520nm | FilterCubeLFOV:0:1:P
96374 | C-FLL LFOV DS RED HISN SEM BL FILTER SET |  | C-FLL Large Field of View DSRed/TRITC/Cy3 Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 554/23nm (542.5-565.5nm), Emission: 609/54nm (582-636nm), Dichroic Mirror: 573nm | FilterCubeLFOV:0:1:P
96375 | C-FLL LFOV TEXAS RED HISN SEM BL SET |  | C-FLL Large Field of View Texas Red/mCherry Filter Set, High-Signal- Noise, Semrock Brightline®, Excitation: 562/40nm (542-582nm), Emission: 641/75nm (603.5-678.5nm), Dichroic Mirror: 593nm | FilterCubeLFOV:0:1:P
96376 | C-FLL LFOV CY5 HISN SEM BL FILTER SET |  | C-FLL Large Field of View Cy5 Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 618/50nm (593-643nm), Emission: 698/70nm (663- 733nm), Dichroic Mirror: 652nm | FilterCubeLFOV:0:1:P
96377 | C-FLL LFOV CY7 HISN SEM BL FILTER SET |  | C-FLL Large Field of View Cy7, Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 708/75nm (670.5-745.5nm), Emission: 809/81nm (768.5-849.5nm). Dichroic Mirror: 757nm | FilterCubeLFOV:0:1:P
96378 | C-FLL LFOV QUAD D/F/T/CY5 LED FILTER SET |  | C-FLL Large-Field of View Quad DAPI/FITC/TRITC/CY5 LED Filter Set for Nikon D-LEDI LED Fluorescence Illumination System | FilterCubeLFOV:0:1:P
96380 | C-FL HC DAPI/FITC |  | CF-L ET DAPI/FITC Dual Band Filter Set | FilterCube:0:1:P
96381 | C-FL HC FITC/TRITC |  | C-FL ET FITC/TRITC Dual Band Filter Set | FilterCube:0:1:P
96382 | C-FL HC FITC/TEXAS RED |  | CF-L ET FITC/Texas Red Dual Band Filter Set | FilterCube:0:1:P
96383 | C-FL HC DAPI/FITC/TRITC |  | CF-L ET DAPI/FITC/TRITC Triple Band Filter Set | FilterCube:0:1:P
96384 | C-FL HC DAPI/FITC/TEXAS RED |  | CF-L ET DAPIFITC/Texas Red Triple Band Filter Set | FilterCube:0:1:P
96389 | C-FL DAPI SOLA HISN SEM BL FILTER SET |  | C-FL DAPI SOLA Filter Set, High-Signal-Noise, Semrock BrightLine®, Excitation: 378/53nm (352-404nm), Emission: 447/60nm (417-477nm), Dichroic Mirror: 409nm | FilterCube:0:1:P
96390 | C-FL DAPI HISN SEM BL FILTER SET |  | C-FL DAPI Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 356/30nm (341-371nm), Emission: 447/60nm (417-477nm), Dichroic Mirror: 409nm | FilterCube:0:1:P
96391 | C-FL CFP HC HISN SEM BL FILTER SET |  | C-FL CFP Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 438/24nm (426-450nm), Emission: 483/32nm (467-499nm), Dichroic Mirror: 458nnm | FilterCube:0:1:P
96392 | C-FL GFP HC HISN SEM BL FILTER SET |  | C-FL GFP/FITC/Cy2 Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 466/40nm (446-486nm), Emission: 525/50nm (500-550nm), Dichroic Mirror: 495nm | FilterCube:0:1:P
96393 | C-FL YFP HC HISN SEM BL FILTER SET |  | C-FL YFP Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 500/24nm (488-515nm), Emission: 542/27nm (528.5-555.5nm), Dichroic Mirror: 520nm | FilterCube:0:1:P
96394 | C-FL DS RED HISN SEM BL FILTER SET |  | C-FL DSRed/TRITC/Cy3 Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 554/23nm (542.5-565.5nm), Emission: 609/54nm (582-636nm), Dichroic Mirror: 573nm | FilterCube:0:1:P
96395 | C-FL TEXAS RED HISN SEM BL FILTER SET |  | C-FL Texas Red/mCherry Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 562/40nm (542-582nm), Emission: 641/75nm (603.5-678.5nm), Dichroic Mirror: 593nm | FilterCube:0:1:P
96396 | C-FL CY5 HISN SEM BL FILTER SET |  | C-FL Cy5 Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 618/50nm (593-643nm), Emission: 698/70nm (663-733nm), Dichroic Mirror: 652nm | FilterCube:0:1:P
96397 | C-FL CY7 HISN SEM BL FILTER SET |  | C-FL Cy7, Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 708/75nm (670.5-745.5nm), Emission: 809/81nm (768.5- 849.5nm). Dichroic Mirror: 757nm | FilterCube:0:1:P
96398 | C-FL QUAD D/F/T/CY5 LED FILTER SET |  | C-FL Quad DAPI/FITC/TRITC/CY5 LED Filter Set for Nikon D-LEDI LED Fluorescence Illumination System | FilterCube:0:1:P
96400 | C-FLL DAPI HC SOLA HISN ZERO SHIFT |  | C-FLL DAPI SOLA Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 395/25nm (382.5-407.5nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 425nm | FilterCube:0:1:P
96401 | C-FLL DAPI HC HISN ZERO SHIFT |  | C-FLL DAPI Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 350/50nm (325-375nm), Emission: 460/50nm (435-485nm), Dichroic Mirror: 400nm | FilterCube:0:1:P
96402 | C-FLL CFP HC HISN ZERO SHIFT |  | C-FLL CFP Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 436/20nm (426-446nm), Emission: 480/40nm (460-500nm), Dichroic Mirror: 455nm | FilterCube:0:1:P
96403 | C-FLL GFP HC HISN ZERO SHIFT |  | C-FLL GFP Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 470/40nm (450-490nm), Emission: 525/50nm (500-550nm), Dichroic Mirror: 495nm | FilterCube:0:1:P
96404 | C-FLL YFP HC HISN ZERO SHIFT |  | C-FLL YFP Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 500/20nm (490-510nm), Emission: 535/30nm (520-550nm), Dichroic Mirror: 515nm | FilterCube:0:1:P
96405 | C-FLL DS RED HC HISN ZERO SHIFT |  | C-FLL DS Red Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 545/30nm (530-560nm), Emission: 620/60nm (590-630nm), Dichroic Mirror: 570nm | FilterCube:0:1:P
96406 | C-FLL TEXAS RED HC HISN ZERO SHIFT |  | C-FLL Texas Red Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 560/40nm (540-580nm), Emission: 630/70nm (593-668nm), Dichroic Mirror: 585nm | FilterCube:0:1:P
96407 | C-FLL CY5 HC HISN ZERO SHIFT |  | C-FLL CY5 Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 620/60nm (590-650nm), Emission: 700/75nm (663-738nm), Dichroic Mirror: 660nm | FilterCube:0:1:P
96408 | C-FLL CY7 HC HISN ZERO SHIFT |  | C-FLL CY7 Hard Coat, High Signal-to-Noise, Zero Shift Filter Set, Excitation: 710/75nm (672.5-747.5nm), Emission: 810/90nm (765-855nm), Dichroic Mirror: 760nm | FilterCube:0:1:P
97240 | C-FL TIRF 488 ARGON FTR (FITC) W/CLEANUP |  | C-Fl TIRF 488 Argon Filter (FITC) with Cleanup | FilterCube:0:1:P
97241 | C-FL TIRF 543G-HENE FTR-TRITC-W/CLEANUP |  | C-Fl TIRF 543G-Hene Filter-TRITC with Cleanup | FilterCube:0:1:P
97242 | C-FL TIRF 643R-HENE FTR (CY5) W/CLEANUP) |  | C-Fl TIRF 643R-Hene Filter (CY5) with Cleanup | FilterCube:0:1:P
97243 | C-FL FITC/TRITC W/O EXCITER |  | C-Fl FITC/TRITC without Exciter | FilterCube:0:1:P
97244 | C-FL FITC/CY3/CY5 W/O EXCITER |  | C-Fl FITC/CY3/CY5 without Exciter | FilterCube:0:1:P
97245 | C-FL TIRF 442 FILTER (CFP) W/CLEANUP |  | C-Fl TIRF 442 Filter (CFP) with Cleanup | FilterCube:0:1:P
97246 | C-FL TIRF 514 FILTER (YFP) W/CLEANUP |  | C-Fl TIRF 514 Filter (YFP) with Cleanup | FilterCube:0:1:P
97253 | C-FL TIRF 56I FILTER TRITC W/ CLEANUP |  | C-Fl TIRF 56I Filter TRITC with Cleanup | FilterCube:0:1:P
97270 | C-FL SRIC FILTER SET |  | C-FL Surface Reflection Interference Contrast (SRIC) Cube, Excitation: 535/80nm (495-575nm), Dichroic Mirror: 400nm, No Emission Filter | FilterCube:0:1:P
97271 | C-FL SRIC FILTER SET 2 |  | C-FL SRIC Filter Set 2 | FilterCube:0:1:P
97273 | C-FL-HQ 488/GFP HI PERFORMANCE TIRF SET |  | C-Fl-Hq 488/Gfp Hi Performance TIRF Set | FilterCube:0:1:P
97319 | C-TIRF ULTRA HI S/N 532NM TIRF SET |  | C-TIRF Ultra High Signal-to-Noise 532nm TIRF Filter Set, 532/10nm HC Cleanup, HC TIRF 432nm Dichroic Mirror in Metal Cube, 640/75nm HC Barrier and Supplementary LP 542nm Laser Block Barrier Filter (4 filters total). | FilterCube:0:1:P
97320 | C-TIRF ULTRA HI S/N 440NM TIRF SET |  | C-TIRF Ultra High Signal-to-Noise 440nm TIRF Filter Set, 442/10nm HC Cleanup, HC TIRF 442nm Dichroic Mirror in Metal Cube, 480/40nm HC Barrier and Supplementary LP 460nm Laser Block Barrier Filter (4 filters total) | FilterCube:0:1:P
97321 | C-TIRF ULTRA HI S/N 488NM TIRF SET |  | C-TIRF Ultra High Signal-to-Noise 488nm TIRF Filter Set, 488/10nm HC Cleanup, HC TIRF 488nm Dichroic Mirror in Metal Cube, 525/50nm HC Barrier and Supplementary LP 500nm Laser Block Barrier Filter (4 filters total) | FilterCube:0:1:P
97322 | C-TIRF ULTRA HI S/N 514NM TIRF SET |  | C-TIRF Ultra High Signal-to-Noise 514nm TIRF Filter Set, 514/10nm HC Cleanup, HC TIRF 514nm Dichroic Mirror in Metal Cube, 540/40nm HC Barrier and Supplementary LP 525nm Laser Block Barrier Filter (4 filters total) | FilterCube:0:1:P
97323 | C-TIRF ULTRA HI S/N 561NM TIRF SET |  | C-TIRF Ultra High Signal-to-Noise 561nm TIRF Filter Set, 561/10nm HC Cleanup, HC TIRF 561nm Dichroic Mirror in Metal Cube, 600/50nm HC Barrier and Supplementary LP 575nm Laser Block Barrier Filter (4 filters total) | FilterCube:0:1:P
97324 | C-TIRF ULTRA HI S/N 638NM TIRF SET |  | C-TIRF Ultra High Signal-to-Noise 638nm TIRF Filter Set, 688/20nm HC Cleanup, HC TIRF 638nm Dichroic Mirror in Metal Cube, 700/75nm HC Barrier and Supplementary LP 660nm Laser Block Barrier Filter (4 filters total) | FilterCube:0:1:P
97325 | C-TIRF ULTRA HI S/N DUAL 442/514NM SET |  | C-TIRF Ultra High Signal-to-Noise 442/514nm CFP/YFP Dual TIRF Filter Set, Dual 442/514nm Laser Cleanup, HC TIRF 442/514nm Dual Dichroic Mirror in Metal Cube, HC Dual Barrier Filter and two Supplementary BP 470/40nm and 540/40nm HC Laser Block Barrier Filters 25mm for Emission Filter Wheel. (5 filters total) | FilterCube:0:1:P
97326 | C-TIRF ULTRA HI S/N DUAL 488/561NM SET |  | C-TIRF Ultra High Signal-to-Noise 488/561nm Blue/Green Dual TIRF Filter Set, Dual 488/561nm Laser Cleanup, HC TIRF 488/561nm Dual Dichroic Mirror in Metal Cube, HC Dual Barrier Filter and two Supplementary BP 525/50nm and 600/50nm HC Laser Block Barrier Filters 25mm for Emission Filter Wheel. (5 filters total) | FilterCube:0:1:P
97327 | C-TIRF ULTRA HI S/N QUAD 405/488/561/638 |  | C-TIRF Ultra High Signal-to-Noise 405/488/561/638nm Quad TIRF Filter Set, Quad 405/488/561/638nm Laser Cleanup, HC TIRF 405/488/561/638nm Quad Dichroic Mirror in Metal Cube, HC Quad Barrier Filter and three Supplementary BP 525/50nm, 600/50nm and 700/75nm HC Laser Block Barrier Filters 25mm for Emission Filter Wheel. (6 filters total) | FilterCube:0:1:P
97329 | C-TIRF ULTRA HI S/N 543NM TIRF SET |  | C-TIRF Ultra High Signal-to-Noise 543nm TIRF Filter Set, 543/10nm HC Cleanup, HC TIRF 543nm Dichroic Mirror in Metal Cube, 600/50nm HC Barrier and Supplementary LP 560nm Laser Block Barrier Filter (4 filters total) | FilterCube:0:1:P
97331 | C-TIRF ULTRA HI S/N DUAL 488/594NM SET |  | C-TIRF Ultra High Signal-to-Noise 488/594nm Green/Red Dual TIRF FilterSet, Dual 488/594nm Laser Cleanup, HC TIRF 488/594nm Dual DichroicMirror in Metal Cube, HC Dual Barrier Filter <(>&<)> two SupplementaryBP525/50nm and 640/50nm HC Laser Block Barrier Filters 25mm for emissionfilter wheel. (5 filters total) | FilterCube:0:1:P
97333 | C-TIRF ULTRA HI S/N TRIPLE 442/514/594NM |  | C-TIRF Ultra High Signal-to-Noise 445/514/594nm CFP/YFP/mCherry TripleTIRF Filter Set, Triple 445/514/594nm Laser Cleanup, HC TIRF445/514/594nm Triple Dichroic Mirror in Metal Cube, HC Triple BarrierFilter <(>&<)> three Supplementary BP 480/40nm, 545/40nm and 650/75nm HCLaserBlock Barrier Filters 25mm for emission filter wheel. (6 filters total) | FilterCube:0:1:P
97335 | C-NSTORM QUAD 405/488/561/647 FILTER SET |  | C-TIRF Ultra High Signal-to-Noise 405/488/561/647nm Quad TIRF Filter Set, Quad 405/488/561/647nm Laser Cleanup, HC TIRF 405/488/561/647nm Quad Dichroic Mirror in Metal Cube, HC Quad Emission Filter and Three Supplementary BP 525/50nm, 600/50nm and 700/72nm HC Laser Block Barrier Filters 25mm for Emission Filter Wheel. (6 filters total) | FilterCube:0:1:P
97336 | C-NSTORM 647 FILTER SET |  | C-N STORM Ultra High Signal-to-Noise 647nm NSTORM Filter Set | FilterCube:0:1:P
97337 | C-TIRF 405 REFLECT 320 & 720 TRANSMIT |  | C-TIRF Filter Cube 405nm Reflection with 320nm and 720nm Transmission | FilterCube:0:1:P
97338 | C-NSTORM QUAD 405/488/561/640 FILTER SET |  | C-NSTORM Ultra High Signal-to-Noise 405/488/561/640nm Quad STORM/TIRF Filter Set, Quad 405/488/561/640nm Laser Cleanup, Quad 405/488/561/640 Polychroic Mirror, Quad 405/488/561/640m Laser Emitter, All Optics Installed in Cube and Laser Aligned | FilterCube:0:1:P
77014398 | ET525/50M EMISSION FILTER, 25MM |  | ET525/50m, Emission filter, 25mm. | 25mmFilter:0:1:P
77014400 | ET700/75M EMISSION FILTER 25MM |  | ET700/75m emission filter, 25mm | 25mmFilter:0:1:P
77014724 | ET450/50M, EMISSION FILTER, 25MM |  | ET450/50M, Emission Filter, 25mm, | 25mmFilter:0:1:P
77015565 | 685NM EMISSION FILTER 25MM |  | 685nm 25mm Brightline Emission Filter, Alexa 647 | 25mmFilter:0:1:P
77015728 | FF02-809/81-25 EMISSION FILTER, 25MM |  | FF02-809/81-25 Bright-Line Band Pass Emission Filter, 25mm, unmounted | 25mmFilter:0:1:P
77015743 | FF02-525/40-25 EMISSION FILTER 25MM |  | FF02-525/40-25 Emission Filter 25mm | 25mmFilter:0:1:P
77015824 | FF02-809/81-32 EMISSION FILTER 32MM |  | FF02-809/81-32 Brightline Emission Filter, 32mm | 32mmFilter:0:1:P
77015839 | FF01-595/31-25 EMISSION FILTER, 25MM |  | FF01-595/31-25 Emission Filter, 25mm | 25mmFilter:0:1:P
77015841 | FF02-641/75-25 SINGLE BAND BP FILTR 25MM |  | FF02-641/75-25 Single-band Bandpass filter, 25mm | 25mmFilter:0:1:P
77015857 | FF01-474/27-32 EMISSION FILTER 32MM |  | FF01-474/27-32 Emission Filter, 32mm | 32mmFilter:0:1:P
77015870 | FF01-432/36 EMISSION FILTER, 25MM |  | FF01-432/36 Brightline Emission Filter, 25mm | 25mmFilter:0:1:P
77015871 | FF01-515/30 EMISSION FILTER, 25MM |  | FF01-515/30 Brightline Emission Filter, 25mm | 25mmFilter:0:1:P
77015882 | FF01-474/27-25 EXCITATION FILTER 25MM |  | FF01-474/27-25 Excitation Filter 25mm | 25mmFilter:0:1:P
77015891 | FF01-680/42-32 EMISSION FILTER 32MM |  | FF01-680/42-32 Emission Filter 32mm | 32mmFilter:0:1:P
77015893 | FF01-544/24-32 EMISSION FILTER 32MM |  | FF01-544/24-32 Emission Filter 32mm | 32mmFilter:0:1:P
77015894 | FF02-641/75-32 EMISSION FILTER 32MM |  | FF02-641/75-32 Emission Filter 32mm | 32mmFilter:0:1:P
77015898 | FF01-544/24-25 EMISSION FILTER 25MM |  | FF01-544/24-25 Emission Filter 25mm | 25mmFilter:0:1:P
77015913 | 560/25NM SINGLE-BAND BP FLTR 25MM ROUND |  | 560/25 nm BrightLine single-band bandpass filter, 25mm round. | 25mmFilter:0:1:P
77015914 | 485/20NM SINGLE-BAND BP FLTR 25MM ROUND |  | 485/20 nm BrightLine single-band bandpass filter, 25mm round. | 25mmFilter:0:1:P
77015915 | 819/44NM SINGLE-BAND BP FLTR 25MM ROUND |  | 819/44 nm BrightLine single-band bandpass filter, 25mm round. | 25mmFilter:0:1:P
77015916 | 511/20NM SINGLE-BAND BP FLTR 25MM ROUND |  | 511/20 nm BrightLine single-band bandpass filter, 25mm round. | 25mmFilter:0:1:P
77015917 | 438/24NM SINGLE-BAND BP FLTR 25MM ROUND |  | 438/24 nm BrightLine single-band bandpass filter, 25mm round. | 25mmFilter:0:1:P
77015918 | POLYCHRIC/EM F/RETRA FURA MTD TI CUBE |  | Filter cube for Lumencor RETRA FURA light source. FF409-Di03-25x36dichroic and FF01-510/84-25 emitter mounted in Ni/Ti cube. No excitationfilter (included in light source). | FilterCubeLFOV:0:1:P
77015919 | PLYCHRC/EM F/RETRA FURA MTD LFOV TI2 CBE |  | LFOV Ti2 filter cube for Lumencor RETRA FURA light source. FF409-Di03-t2-32x44-Ti2 dichroic andFF01-510/84-32 emitter mounted in Ti2 LFOVcube. No excitation filter (included in light source). | FilterCubeLFOV:0:1:P
77015924 | LED-DA/FI/TR/Cy5/Cy7 FLTR ST MTD TI CB |  | Penta band filter cube for LED light sources LED-DA/FI/TR/Cy5/Cy7,mounted in standard Ni/Ti cube.Exciter: FF01378/474/554/635/735-25Emitter: FF01432/515/595/681/809-25Polychroic: FF409/493/573/652/759Di01-25x36 | FilterCube:0:1:P
77015925 | LED-DA/FI/TR/Cy5/Cy7 FLTRST MT TI2 LFOV |  | LFOV Ti2 penta band filter cube for LED light sources LED-DA/FI/TR/Cy5/Cy7, mounted in Ti2 LFOV cube.Exciter: FF01378/474/554/635/735-25Emitter: FF01432/515/595/681/809-25Polychroic: FF409/493/573/652/759Di01-25x36 | FilterCubeLFOV:0:1:P
77015944 | CUSTOM QUAD SET DICHROIC/EM MTD TI CUBE |  | Custom Filter set mounted Ti Cube.Dichroic:FF409/493/573/652-Di02-25x36Emitter: FF01-432/515/595/730-25 | FilterCube:0:1:P
77015964 | FF01-515/30-32 EMISSION FILTER 32MM |  | FF01-515/30-32 Emission Filter 32mm | 32mmFilter:0:1:P
77015965 | FF01-595/31-32 EMiSSION FILTER 32MM |  | FF01-595/31-32 Emission Filter 32mm | 32mmFilter:0:1:P
77015970 | SPECTRAIII 390/475/555/635/747 MTD TI CB |  | Custom Spectra III set (390/475/555/635/747) mounted in Ti cubePolychroic: FF-409/493/573/652/759-Di01-25x36Emitter: FF01-432/515/595/681/809-25 | FilterCube:0:1:P
77015971 | SPCTRIII 390/475/555/635/747 MT TI2 LFOV |  | Custom Spectra III set (390/475/555/635/747) mounted in Ti2 LFOV cubePolychroic: FF-409/493/573/652/759-Di01-32x44Emitter: FF01-432/515/595/681/809-32 | FilterCubeLFOV:0:1:P
77015972 | SPECTRA III 440/510/575 MTD TI CUBE |  | Custom Spectra III set (440/510/575) mounted in Ti cubePolychroic: FF-459/526/596-Di01-25x36Emitter: FF01-475/543/702-25 | FilterCube:0:1:P
77015973 | SPECTRA III 440/510/575 MTD TI2 LFOV CBE |  | Custom Spectra III set (440/510/575) mounted in Ti2 LFOV cubePolychroic: FF-459/526/596-Di01-32x44Emitter: FF01-475/543/702-32 | FilterCubeLFOV:0:1:P
77015974 | FF01-680/42-25 SINGLE BD BP FILTER 25MM |  | FF01-680/42-25 Single Band Band Pass Filter, 25mm | 25mmFilter:0:1:P
77015975 | FF01-432/36-32 SINGLE BD BP FILTER 32MM |  | FF01-432/36-32 Single Band Band Pass Filter, 32mm | 32mmFilter:0:1:P
77015987 | SPECTRA III 390/475/575 MTD TI CB |  | Custom Spectra III set (390/475/575) mounted in Ti cubePolychroic: FF-409/493/596-Di02-25x36Emitter: FF01-432/523/702--25 | FilterCube:0:1:P
77015988 | SPECTRA III 390/475/575 MTD TI2 LFOV |  | Custom Spectra III set cube (390/475/575) mounted in Ti2 LFOVPolychroic: FF-409/493/596-Di02-32x44Emitter: FF01-432/523/702--32 | FilterCubeLFOV:0:1:P
77015989 | FF02-525/40-32 SINGLE BP FILTER 32 MM |  | FF02-525/40-32 Single Bandpass Filter 32 mm | 32mmFilter:0:1:P
77074328 | ET595/40M EM FLTR 25MM RED F/SPECTRAX |  | ET595/40m Emission Filter, Red Emission for SpectraX quad band, 25mm | 25mmFilter:0:1:P
77074590 | ET600/50M EMISSION FILTER 25MM |  | ET600/50m Emission Filter 25mm | 25mmFilter:0:1:P
77074693 | TRF89901EMV2 ET405/488/561/640NM TIRF CB |  | "TRF89901-EMv2 - ET405/488/561/640nm Laser Quad Band Set for TIRFapplications mounted in a TIRF cubeExciter: ZET405/488/561/640xv2Polychroic: ZT405/488/561/640rpcv2Emitter: ET445/58mEmitter: ET525/50mEmitter: ET595/44mEmitter: ET706/95mEmitter: ZET405/488/561/640mv2" | FilterCube:0:1:P
77074977 | 79001 ET FURA2 FLTR ST MTED TI CUBE |  | 79001 ET Fura2 Filter Set, mounted TE2000/Ti cube.Exciter: ET340x, 25mmExciter: ET380x, 25mmDichroic: T400LP, 25.5x36x2mmEmission: ET510/80m, 25mm | FilterCube:0:1:P
77075018 | ZET408/473/545/635/750M EM FILTER 25MM |  | ZET408/473/545/635/750m Penta Emission Filter for Cicero, 25mm | 25mmFilter:0:1:P
77075024 | 89402M-OD8 QUADBAND EM 25MMRX5MM RING |  | 89402m-OD8 Quadband Emission Filter 25mmR x 5mm ring
77081015 | DAPI/FITC/CY3/CY5 DUAL CAM F/CRST/CELSTA |  | DAPI/FITC/CY3/CY5 dual camera filter set for Crest V3 with Celesta.Includes dual pass dichroic 32x44x2mm, D/CY3 and F/CY5 emitters in 32mm. | 25mmFilter:0:1:P
77081023 | FF01-794/32-25 SINGLE BP FILTER 25MM |  | FF01-794/32-25 Single Band Pass Filter, 25mm, 3.5mm | 25mmFilter:0:1:P
77081088 | CUSTOM QUAD SET DICHROIC/EM MT TI2 FLOV |  | Custom Filter set mounted Ti2 Large FOV cube.Dichroic:FF409/493/573/652-Di02-32x44-Ti2Emitter: FF01-432/515/595/730-32 | FilterCubeLFOV:0:1:P
77196370 | C-FLL LFOV BFP HISN SEM BL FILTER |  | C-FLL Large Field of View BFP Filter Set, High-Signal-Noise, Semrock Brightline®, Excitation: 356/30nm (341-371nm), Emission: 447/60nm (417- 477nm), Dichroic Mirror: 409nm | FilterCubeLFOV:0:1:P
MBE41201 | C-FL UV-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL UV-2A for LV-UEPI2 and D-Fl U-Epi. EX: 330~380nm, DM:400nm,BA:420nm | FilterCube:0:1:P
MBE41204 | C-FL-C UV-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL-C UV-2A for LV-UEPI2 and D-FL U-EPI. EX: 330~380nm, DM:400nm,BA:420nm | FilterCube:0:1:P
MBE42101 | C-FL V-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL V-2A for LV-UEPI2 and D-Fl U-Epi | FilterCube:0:1:P
MBE42104 | C-FL-C V-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL-C V-2A for LV-UEPI2 and D-Fl U-Epi | FilterCube:0:1:P
MBE43100 | C-FL BV-2A |  | "C-FL BV-2A, Excitation: 400-440nm, Emission: 470nm, Dichroic Mirror: 455nm." | FilterCube:0:1:P
MBE43101 | C-FL BV-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL BV-2A for LV-UEPI2 and D-Fl U-EPI. EX: 400~440nm, DM: 455nm,BA:470nm | FilterCube:0:1:P
MBE43104 | C-FL-C BV-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL-C BV-2A for LV-UEPI2 and D-FL U-EPI. EX: 400~440nm, DM: 455nm,BA:470nm | FilterCube:0:1:P
MBE43113 | C-FLL-C CFP HQ L FOV |  | C-FLL-C CFP HQ | FilterCubeLFOV:0:1:P
MBE44200 | * C-FL EPI-FL FILTER MODULE N B-3A (LTD) |  | * C-FL EPI-FL FILTER MODULE N B-3A (LTD) | FilterCube:0:1:P
MBE44501 | C-FL B-2A FOR LV-UEPI2 AND D-FL U-EPI |  | B-2A, Excitation: 450-490nm, Emission: 520nm Longpass, Dichroic Mirror: 505nm | FilterCube:0:1:P
MBE44504 | C-FL-C B-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL-C B-2A for LV-UEPI2 and D-FL U-EPI. EX: 450~490nm, DM: 505nm,BA:520nm | FilterCube:0:1:P
MBE44723 | C-FLL-C FITC L FOV |  | C-FLL-C FITC | FilterCubeLFOV:0:1:P
MBE44730 | C-FL EPI-FL FILTER GFP-LP |  | C-FL EPI-FL Filter GFP-LP | FilterCube:0:1:P
MBE44903 | C-FLL-C YFP HQ L FOV |  | C-FLL-C YFP HQ | FilterCubeLFOV:0:1:P
MBE45501 | C-FL G-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL G-2A for LV-UEPI2 and D-Fl U-EPI. EX: 510~560nm, DM:575nm,BA:590nm | FilterCube:0:1:P
MBE45504 | C-FL-C G-2A FOR LV-UEPI2 AND D-FL U-EPI |  | C-FL-C G-2A for LV-UEPI2 and D-FL U-EPI. EX: 510~560nm, DM:575nm,BA:590nm | FilterCube:0:1:P
MBE45600 | C-FL EPI-FL FILTER TRITC |  | C-FL EPI-FL Filter TRITC | FilterCube:0:1:P
MBE45703 | C-FLL-C CY3 L FOV |  | C-FLL-C CY3 | FilterCubeLFOV:0:1:P
MBE47100 | C-FL-HQ FILTER BLOCK |  | C-FL-HQ Filter Block | FilterCube:0:1:P
MBE49300 | C-FLF FRET CUBE |  | C-FLF FRET Cube | FilterCube:0:1:P
MBE49310 | C-FLM MIRROR CUBE |  | C-FLM Mirror Cube | FilterCube:0:1:P
MXA22128 | C-FL2 BLANK CUBE |  | C-FL2 Blank Cube | FilterCube:0:1:P
77014027 | POLARIZING FILTER 25MM |  |  | 25mmFilter:0:1:P
77014465 | 49006 ET CY5 FILTER SET 25MM UNMTED |  |  | 25mmFilter:0:1:P
77014536 | ET470/24M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77014537 | ET535/30M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77014539 | ET632/60M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77014544 | ET480/40M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77014571 | ET470/40X GFP EXCITATION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77014654 | ET605/70M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77014668 | 69008M CFP/YFP/MCHERRY ET EM FILTR 25MM |  |  | 25mmFilter:0:1:P
77014694 | ET540/30M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77014812 | ET585/40M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77015520 | FF01-593/40-25 EM FILTER 25MM |  |  | 25mmFilter:0:1:P
77015532 | FF01-472/30-25 CFP EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77015540 | FF01-624/40-25 TX RD EMISSION FILTER 25M |  |  | 25mmFilter:0:1:P
77015716 | 664NM EDGEBASIC LP FILTER, 25MM |  |  | 25mmFilter:0:1:P
77015831 | FF01-708/75-25 SINGLE-BAND BP FILTR 25MM |  |  | 25mmFilter:0:1:P
77015905 | 391/477/549/638.5/741 PENTA EX 25 MM |  |  | 25mmFilter:0:1:P
77015911 | FF01-593/LP-25 LP FILTER 25MM |  |  | 25mmFilter:0:1:P
77015941 | FF01-640/40-25 SINGLE BAND FILTER 25MM |  |  | 25mmFilter:0:1:P
77015945 | 568NM EDGEBASIC LP FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074014 | 79001 ET FURA2 FLTR ST, 32MM MTD TI CUBE |  |  | FilterCube:0:1:P
77074266 | 59217 ET CFP/YFP SET, 32MM MTED TI CUBE |  |  | FilterCube:0:1:P
77074322 | TRF89901-EM-ET QUAD TIRF ST MTD TIRF CUB |  |  | FilterCube:0:1:P
77074326 | ET435/26M EM FLTR 25MM BLUE F/SPECTRAX |  |  | 25mmFilter:0:1:P
77074327 | ET510/40M EM FLTR 25MM GREEN F/SPECTRAX |  |  | 25mmFilter:0:1:P
77074379 | CUSTOM 405/488/561/638 EX/DI MTD TIRF CB |  |  | FilterCube:0:1:P
77074402 | ET545/40M LONG PASS EMISSION FILTER, 25M |  |  | 25mmFilter:0:1:P
77074416 | ET645/75M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074420 | ET700/50M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074442 | ET680/30M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074444 | ZET640/20X EXCITATION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074497 | AT635/60M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074519 | ET485/25M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074538 | ET610LP LONG PASS EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074539 | ET655LP LONG PASS EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074543 | ET590/50M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074580 | ND 0.2 B, NEUTRAL DENSITY FILTER 32MM |  |  | 32mmFilter:0:1:P
77074583 | ND 1.0 B, NEUTRAL DENSITY FILTER 32MM |  |  | 32mmFilter:0:1:P
77074587 | ND 3.0 B, NEUTRAL DENSITY FILTER 32MM |  |  | 32mmFilter:0:1:P
77074594 | 49002 ET GFP/CY2 FLTR SET MTED TIRF CUBE |  |  | FilterCube:0:1:P
77074608 | 79006 INDO FILTER SET MTD TI CUBE |  |  | FilterCube:0:1:P
77074648 | ZET532/640M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074658 | SPECTRA X DAPI/FITC/CY3/CY5 EM SET 32MM |  |  | 32mmFilter:0:1:P
77074664 | 69002M DAPI/FITC/TX RD EM FILTER 25MM |  |  | 25mmFilter:0:1:P
77074682 | ET700LP LONG PASS EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074689 | ET600/40M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074692 | ZET405/488/561/640MV2 QUAD EM FLTR 25MM |  |  | 25mmFilter:0:1:P
77074705 | ET585/65M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074729 | ET572/23M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074759 | ET510/20M EMSSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074768 | ET525/50M EMISSION FILTER 32MM |  |  | 32mmFilter:0:1:P
77074769 | ET600/50M EMISSION FILTER 32MM |  |  | 32mmFilter:0:1:P
77074771 | ZT405/488/561/640RPCV2 MTD TIRF CUBE. |  |  | FilterCube:0:1:P
77074780 | 33002v2 FLTR SET MTD TI CUBE |  |  | FilterCube:0:1:P
77074793 | ET615/40M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074803 | ZET405/488M CUSTOM DUAL EM FILTER 25MM |  |  | 25mmFilter:0:1:P
77074804 | ZET561/640M DUAL EM FILTER CY3/CY5 25MM |  |  | 25mmFilter:0:1:P
77074822 | ZT442DCRB-UF2 MOUNTED IN TIRF CUBE. |  |  | FilterCube:0:1:P
77074823 | ZT488DCRB-UF2 MOUNTED IN TIRF CUBE. |  |  | FilterCube:0:1:P
77074824 | ZT518DCRB-UF2 MOUNTED IN TIRF CUBE. |  |  | FilterCube:0:1:P
77074825 | ZT561DCRB-UF2 MOUNTED IN TIRF CUBE. |  |  | FilterCube:0:1:P
77074841 | ET610LP LONGPASS FILTER 25MM |  |  | 25mmFilter:0:1:P
91174 | 473NM LASER REFL MIRROR MTD TI2 LFOV |  |  | FilterCubeLFOV:0:1:P
77074848 | 49309 ET ORANGE #2 FISH MTD TI2 LFOV CBE |  |  | FilterCubeLFOV:0:1:P
77074849 | 49310 ET RED #2 FISH MTD TI2 LFOV CUBE |  |  | FilterCubeLFOV:0:1:P
77081013 | PENTABD FLTR ST F/LFOV CRST/CELSTA UNMTD |  |  | FilterCubeLFOV:0:1:P
77081014 | DUAL BD FLTR ST F/LFOV CRST/CELSTA UNMTD |  |  | FilterCubeLFOV:0:1:P
77015807 | FF03-575/25-25 BP FILTER 25MM |  |  | 25mmFilter:0:1:P
77074915 | ET595/44M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074922 | ET655LP EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074981 | ET445/58M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77074990 | ET667/30M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77075026 | ET706/95M EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77081043 | FF01-565/24-25 SINGLE BP FILTER 25MM |  |  | 25mmFilter:0:1:P
77081070 | FF01-676/37-25 SINGLE BP FILTER 25MM |  |  | 25mmFilter:0:1:P
77081071 | FF01-732/68-25 SINGLE BP FILTER 25MM |  |  | 25mmFilter:0:1:P
77081123 | FF01-449/520-25 DUAL-BD FILTER 25MM |  |  | 25mmFilter:0:1:P
77075011 | ET470/24M EMISSION FILTER 32MM |  |  | 32mmFilter:0:1:P
77081001 | FF01-615/24-32 SINGLE BP FILTER 32MM |  |  | 32mmFilter:0:1:P
77074422 | CUSTOM 647 FILTER SET MTED TIRF CUBE |  |  | FilterCube:0:1:P
77074914 | 89904ET 405/445/514/561/640NM MT TIRF CB |  |  | FilterCube:0:1:P
77074963 | CUSTOM 488/532/640 TIRF SET MT TIRF CUBE |  |  | FilterCube:0:1:P
77074974 | CUSTOM 405 514 SET MTD TIRF CUBE |  |  | FilterCube:0:1:P
77074521 | T387LP DICHROIC MOUNTED IN TI CUBE |  |  | FilterCube:0:1:P
77074522 | T455LP DICHROIC MOUNTED IN TI CUBE |  |  | FilterCube:0:1:P
77074937 | CUSTOM DAPI LP 385NM FLTR ST MTD TI CUBE |  |  | FilterCube:0:1:P
77081090 | CY5.5 ZERO SHFT FILTER SET MTED TI CUBE |  |  | FilterCube:0:1:P
77014567 | ET455/50M, DAPI-ET EMISSION FILTER 25MM |  |  | 25mmFilter:0:1:P
77014803 | ET525/36M EMISSION FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074493 | ET605/52M EMISSION FLTER 25MM |  |  | 25mmFilter:0:1:P
77074329 | ET705/72M EM FLTR 25MM FAR RED F/SPCTRAX |  |  | 25mmFilter:0:1:P
77074020 | 59004M FITC/TRITC DUAL BP FLTR, 25MM |  |  | 25mmFilter:0:1:P
77074021 | 69000M DAPI/FITC/TRITC TRPL BP FLTR 25MM |  |  | 25mmFilter:0:1:P
77014759 | 89101M EM FLTR DAPI/FITC/TRITC/CY5, 25MM |  |  | 25mmFilter:0:1:P
77014006 | 0.6 ND A, NEUTRAL DENSITY FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074240 | 1.0 ND A, NEUTRAL DENSITY FILTER 25MM |  |  | 25mmFilter:0:1:P
77074238 | 0.2 ND A, NEUTRAL DENSITY FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074241 | 1.3 ND A, NEUTRAL DENSITY FILTER, 25MM |  |  | 25mmFilter:0:1:P
77071849 | ND 3.0 A, NEUTRAL DENSITY FILTER, 25MM |  |  | 25mmFilter:0:1:P
77074517 | ND 2.0 A NEUTRAL DENSITY FILTER, 25MM |  |  | 25mmFilter:0:1:P
MEN51980 | TI-A ANALYZER BLOCK | Analyzer |  | FilterCube:1:1:P
MEN51985 | TI2-C-DICACL ANALYZER CUBE L FOV | Analyzer |  | FilterCubeLFOV:1:1:P

## CFI60
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
Achromat Objectives
MRP00042 | CFI ACHRO FLAT FIELD 4X-NC |  |  | CFI60Objective:0:1:P
MRP00102 | CFI ACHRO FLAT FIELD 10X-NC |  |  | CFI60Objective:0:1:P
MRP00202 | CFI ACHRO FLAT FIELD 20X WD |  |  | CFI60Objective:0:1:P
MRP00402 | CFI ACHRO FLAT FIELD 40X-NC |  |  | CFI60Objective:0:1:P
MRP05422 | CFI ACHR LWD 40XC NA0.55 WD 2.7-1.7MM-NC |  |  | CFI60Objective:0:1:P
MRP00602 | CFI ACHRO FLAT FIELD 60X-NC |  |  | CFI60Objective:0:1:P
MRP04600 | CFI ACHRO FLAT FIELD 60X NCG-NC |  | Note: 0.17mm coverglass is pre-installed on front of objective so that this product can be used as a NCG objective. If the coverglass is damaged, warranty is voided and no repair is posssible. | CFI60Objective:0:1:P
MRP01902 | CFI ACHRO FLAT FIELD 100X OIL-NC |  |  | CFI60Objective:0:1:P
MRP02902 | CFI ACHRO 100X W/IRIS-NC |  |  | CFI60Objective:0:1:P
MRP40102 | CFI ACHR ADL 10X PH1 NA 0.25 WD 6.2MM-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRP46202 | CFI ACR LWD ADL 20X PH1 NA0.4 WD3.0MM-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRP46402 | CFI ACR LWD ADL 40X PH1 NA.55WD 2.1MM-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRP45422 | CFI ACHRO LWD ADL 40XC PH2 NA 0.55-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRP90100 | CFI ACHROMAT 10X BM-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRP20102 | CFI ACHRO FF LD DL 10X NA 0.25 WD 7.0-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRP20202 | CFI ACHRO FLAT FIELD DL 20X-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRP20402 | CFI ACHRO FLAT FIELD DL 40X-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRP21902 | CFI ACHRO FLAT FIELD DL 100X OIL-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH3Module:1:1:H
MRP26202 | CFI ACHRO DL 20X F NA 0.4 WD 3.1 MM-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRP25422 | CFI ACR LWD DL 40XC NA.55 WD1.7-2.7MM-NC |  | Phase Contrast = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
E Plan Achromat Objectives
MRP70040 | CFI E PLAN ACHROMAT 4X-NC |  |  | CFI60Objective:0:1:P
MRP70100 | CFI E PLAN ACHROMAT 10X-NC |  |  | CFI60Objective:0:1:P
MRP70400 | CFI E PLAN ACHROMAT 40X-NC |  |  | CFI60Objective:0:1:P
MRP70600 | CFI E PLAN ACHROMAT 60X OBJECTIVE |  |  | CFI60Objective:0:1:P
MRP71900 | CFI E PLAN ACHROMAT 100X OIL -NC |  |  | CFI60Objective:0:1:P
Plan Achromat Objectives
MRL00012 | CFI PLAN ACHRO UW 1X NA 0.04 WD 3.2MM-NC |  |  | CFI60Objective:0:1:P
MRL00022 | CFI PLAN ACHRO UW 2X NA 0.06 WD 7.5MM-NC |  |  | CFI60Objective:0:1:P
MRL00042 | CFI PLAN ACHRO 4X NA 0.1 WD 30MM-NC |  |  | CFI60Objective:0:1:P
MRL00102 | CFI PLAN ACHRO 10X NA 0.25 WD 10.5MM-NC |  |  | CFI60Objective:0:1:P
MXA20744 | CFI OBJECTIVE ND FILTER |  |  | CFI60Objective:0:1:P
MRL00202 | CFI PLAN ACHRO 20X NA 0.4 WD 1.3MM-NC |  |  | CFI60Objective:0:1:P
MRL00402 | CFI PLAN ACHRO 40X NA 0.65 WD 0.56MM-NC |  |  | CFI60Objective:0:1:P
MRL01502 | CFI PLAN ACHRO 50X OIL NA 0.9 WD 0.35MM- |  |  | CFI60Objective:0:1:P
MRL01903 | CFI PLAN ACHR 100XOIL NA1.25 WD 0.2MM-NC |  |  | CFI60Objective:0:1:P
MRL65900 | CFI PLAN LWD IMSI 100X W/CORR COLLAR |  |  | CFI60Objective:0:1:P | DIC100XI:0:1:H
MRL20102 | CFI PLAN ACHR DL 10X NA0.25 WD 10.5MM-NC |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRL20202 | CFI PLAN ACHRO DL 20X NA 0.4 WD 1.3MM-NC |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRL20402 | CFI PLAN ACHR DL 40X NA0.65 WD0.57 MM-NC |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRL21903 | CFI PL ACHR DL100XOIL NA1.25 WD0.17MM-NC |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH3Module:1:1:H
Plan Fluorite Objectives
MRH00041 | CFI PLAN FLUOR 4X NA 0.13 WD 17.2MM |  |  | CFI60Objective:0:1:P
MRH00045 | CFI PLAN FLUOR 4X NA 0.13 WD 17.2MM CH |  |  | CFI60Objective:0:1:P
MRH00101 | CFI PLAN FLUOR 10X NA 0.3 WD 16MM |  |  | CFI60Objective:0:1:P | DIC10X:0:1:H
MRH00105 | CFI PLAN FLUOR 10X NA 0.3 WD 16MM CH |  |  | CFI60Objective:0:1:P | DIC10X:0:1:H
MRH00201 | CFI PLAN FLUOR 20X NA 0.5 WD 2.1MM |  |  | CFI60Objective:0:1:P | DIC20X¦DIC20X-C:0:1:H
MRH00205 | CFI PLAN FLUOR 20X NA 0.5 WD 2.1MM CH |  |  | CFI60Objective:0:1:P | DIC20X¦DIC20X-C:0:1:H
MRH07240 | CFI PLAN FLUOR 20X MI NA 0.75 WD 0.35 |  | Older version - Limited quantiy of demo stock only | CFI60Objective:0:1:P
MRH07241 | CFI PLAN FLUOR 20X MI NA 0.75 WD 0.35MM |  |  | CFI60Objective:0:1:P | DIC20X¦DIC20X-C:0:1:H
MRH00405 | CFI PLAN FLUOR 40X NA 0.75 WD 0.66MM CH |  |  | CFI60Objective:0:1:P | DIC40XI¦DIC40X-C:0:1:H
MRH00401 | CFI Plan Fluor 40X NA 0.75 0.66mm WD |  |  | CFI60Objective:0:1:P | DIC40XI¦DIC40X-C:0:1:H
MRH01401 | CFI PLAN FLUOR 40X OIL NA1.30 WD 0.20MM |  |  | CFI60Objective:0:1:P | DIC40XII:0:1:H
MRH00602 | CFI PLAN FLUOR 60XA NA 0.85 WD 0.30MM |  |  | CFI60Objective:0:1:P | DIC60XII:0:1:H
MRH02601 | CFI PLAN FLUOR 60X OIL IRIS |  |  | CFI60Objective:0:1:P
MRH01902 | CFI PLAN FLUOR 100X OIL NA 1.3 WD 0.16MM |  |  | CFI60Objective:0:1:P | DIC100XII¦DIC100XII-R:0:1:H
MRH02902 | CFI PLAN FLUOR 100X OIL W/IRIS NA0.5-1.3 |  |  | CFI60Objective:0:1:P | DIC100XII¦DIC100XII-R:0:1:H
MRH20041 | CFI PLAN FLUOR DL 4X NA 0.13 WD 16.5MM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PHLModule:1:1:H
MRH20045 | CFI PLAN FLUOR DL 4X NA 0.13 WD 16.5MM C |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PHLModule:1:1:H
MRH20101 | CFI PLAN FLUOR DL 10X NA 0.3 WD 15.2MM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRH20105 | CFI PLAN FLUOR DL 10X NA 0.3 WD 15.2MM C |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRH10101 | CFI PLAN FLUOR DLL 10X NA 0.3 WD 16MM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRH10105 | CFI PLAN FLUOR DLL 10X NA 0.3 WD 16MM CH |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRH10201 | CFI PLAN FLUOR DLL 20X NA 0.5 WD 2.1MM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRH10205 | CFI PLAN FLUOR DLL 20X NA 0.5 WD 2.1MM C |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRH10401 | CFI PLAN FLUOR DLL 40X NA 0.75 WD 0.66MM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRH10405 | CFI PLAN FLUOR DLL 40X NA 0.75 WD 0.66MM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRH99420 | CFI PLAN FLUOR BM 40X AS |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRH41902 | CFI PLAN FLUOR ADH 100X OIL NA 1.3 |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH3Module:1:1:H
MRH11902 | CFI PLAN FLUOR DLL 100X OIL NA1.3 WD0.16 |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH3Module:1:1:H
Super Fluor & Super Plan Fluor LWD & ELWD Objectives
MRF00040 | CFI SUPER FLUOR 4X NA 0.20 WD 15.5MM |  |  | CFI60Objective:0:1:P
MRF00041 | CFI SUPER FLUOR 4X NA 0.20 WD 15.5MM |  |  | CFI60Objective:0:1:P
MRF00100 | CFI SUPER FLUOR 10X NA 0.5 WD 1.2MM |  |  | CFI60Objective:0:1:P | DIC10X:0:1:H
MRF00101 | CFI SUPER FLUOR 10X NA 0.5 WD 1.1MM |  |  | CFI60Objective:0:1:P | DIC10X:0:1:H
MRF00200 | CFI SUPER FLUOR 20X NA 0.75 WD 1.0MM |  |  | CFI60Objective:0:1:P | DIC20X¦DIC20XI-C:0:1:H
MRF00201 | CFI SUPER FLUOR 20X NA 0.75 WD 1.0MM |  |  | CFI60Objective:0:1:P | DIC20X¦DIC20XI-C:0:1:H
MRF00400 | CFI SUPER FLUOR 40X NA 0.90, 0.30mm (0.26-0.34) WD, Corr 0.11-0.23 |  |  | CFI60Objective:0:1:P | DIC40XI¦DIC40XI-C:0:1:H
MRF01400 | CFI SUPER FLUOR 40X OIL NA 1.3 WD 0.22MM |  |  | CFI60Objective:0:1:P | DIC40XII:0:1:H
MRF01401 | CFI SUPER FLUOR 40X OIL NA 1.3 WD 0.22MM |  |  | CFI60Objective:0:1:P | DIC40XII:0:1:H
MRF02900 | CFI SUPER FLUOR 100X OIL NA1.3 WD 0.2MM |  |  | CFI60Objective:0:1:P
MRH08250 | CFI SUPER PLAN FLUOR LWD 20XC |  |  | CFI60Objective:0:1:P | DIC20XII:0:1:H
MRH08230 | CFI SUPER PLAN FLUOR ELWD 20XC |  |  | CFI60Objective:0:1:P | DIC20XII:0:1:H
MRH08430 | CFI SUPER PLAN FLUOR ELWD 40XC |  |  | CFI60Objective:0:1:P | DIC40XIV:0:1:H
MRH08630 | CFI SUPER PLAN FLUOR ELWD 60XC |  |  | CFI60Objective:0:1:P | DIC60XIII:0:1:H
MRH48250 | CFI SUPER PLAN FLUOR LWD 20XC ADM |  | Phase Contrast Objective = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRH48430 | CFI SUPER PLAN FLUOR ELWD 40XC ADM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRH48230 | CFI SUPER PLAN FLUOR ELWD 20XC ADM |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH1Module:1:1:H
MRH48600 | CFI SUPER PLAN FLUOR ELWD 60XC ADL |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
Plan Apo Lambda Objectives
MRD00025 | CFI PLAN APO LAMBDA 2X |  | Discontinued - Limited Availability | CFI60Objective:0:1:P
MRD00045 | CFI PLAN APO LAMBDA 4X |  | Discontinued - Limited Availability | CFI60Objective:0:1:P
MRD00105 | CFI PLAN APO LAMBDA 10X |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC10X:0:1:H
MRD00205 | CFI PLAN APO LAMBDA 20X |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC20X¦DIC20XI-C:0:1:H
MRD00405 | CFI PLAN APO LAMBDA 40X |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC40XI¦DIC40XI-C:0:1:H
MRD00605 | CFI PLAN APO LAMBDA 60X DRY |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC60XI¦DIC60XI-R:0:1:H
MRD01605 | CFI PLAN APO LAMBDA 60X OIL |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC60XII¦DIC60XII-R:0:1:H
MRD01905 | CFI PLAN APO LAMBDA 100X OIL |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC100XI¦DIC100XI-R:0:1:H
MRD30205 | CFI PLAN APO LAMBDA DM 20X |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRD30405 | CFI PLAN APO LAMBDA DM 40X |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRD30605 | CFI PLAN APO LAMBDA DM 60X DRY |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH2Module:1:1:H
MRD31605 | CFI PLAN APO LAMBDA DM 60X OIL |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH3Module:1:1:H
MRD31905 | CFI PLAN APO LAMBDA DM 100X OIL |  | Phase Contrast Objectives = Green | CFI60Objective:0:1:P | PH3Module:1:1:H
MRD30201 | CFI PLAN APO DM 20X NA 0.75 WD 1MM SPRG |  |  | CFI60Objective:0:1:P
MRD31400 | CFI APO DM 40X OIL NA1.0 WD 0.16MM SPRG |  |  | CFI60Objective:0:1:P
MRD31602 | CFI PLAN APO DM 60X OIL |  |  | CFI60Objective:0:1:P
Plan Apo Lambda D Objectives
MRD70020 | CFI60 PLAN APOCHROMAT LAMBDA D 2X |  |  | CFI60Objective:0:1:P
MRD70040 | CFI60 PLAN APOCHROMAT LAMBDA D 4X |  |  | CFI60Objective:0:1:P
MRD70170 | CFI60 PLAN APOCHROMAT LAMBDA D 10X |  |  | CFI60Objective:0:1:P | DIC10X:0:1:H
MRD70270 | CFI60 PLAN APOCHROMAT LAMBDA D 20X |  |  | CFI60Objective:0:1:P | DIC20XIII:0:1:H
MRD70470 | CFI60 PLAN APOCHROMAT LAMBDA D 40X |  |  | CFI60Objective:0:1:P | DIC40XI¦DIC40XI-C:0:1:H
MRD71670 | CFI60 PLAN APOCHROMAT LAMBDA D 60X Oil |  |  | CFI60Objective:0:1:P | DIC60XIV¦DIC60XIV-R:0:1:H
MRD71970 | CFI60 PLAN APOCHROMAT LAMBDA D 100X Oil |  |  | CFI60Objective:0:1:P | DIC100XII¦DIC100XII-R:0:1:H
Plan Apo VC Objectives
MRD70200 | CFI PLAN APO VC 20X NA 0.75 WD 1MM |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC20X¦DIC20XI-C:0:1:H
MRD07602 | CFI PLAN APO VC 60X WI NA1.2 WD 0.27MM |  |  | CFI60Objective:0:1:P | DIC60XIV¦DIC60XIV-R:0:1:H
MRD01602 | CFI PLAN APO VC 60X OIL NA 1.4 WD 0.13MM |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC60XI¦DIC60XI-R:0:1:H
MRD01901 | CFI PLAN APO VC 100X OIL NA 1.4 WD 0.13 |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC100XI¦DIC100XI-R:0:1:H
Lambda S Objectives
MRD70100 | CFI PLAN APOCHROMAT LAMBDA S 10X |  | Discontinued - Limited Availability | CFI60Objective:0:1:P | DIC10X:0:1:H
MRD77200 | CFI APO LWD 20X WI NA 0.95 LAMBDA S |  |  | CFI60Objective:0:1:P | DIC20XIII:0:1:H
MRD77205 | CFI PLAN APO LWD 20XC WI NA 0.95 LAMBDA S |  |  | CFI60Objective:0:1:P | DIC20XIV:0:1:H
MRD77400 | CFI APO 40X WI 1.25 NA 0.18MM WD LAMBDA |  |  | CFI60Objective:0:1:P | DIC40XII:0:1:H
MRD77410 | CFI APO LWD 40X WI 1.15 NA LAMBDA S |  |  | CFI60Objective:0:1:P | DIC40XI¦DIC40XI-C:0:1:H
MRD77415 | CFI PLAN APO LWD 40X WI 1.15 NA LAMBDA S |  |  | CFI60Objective:0:1:P | DIC40XI¦DIC40XI-C:0:1:H
MRD71600 | CFI APO 60X OIL 1.4NA 0.14MM WD LAMBDA S |  |  | CFI60Objective:0:1:P | DIC60XII¦DIC60XII-R:0:1:H
Silicone Immersion
MRD73250 | CFI PLAN APO LAMBDA S 25XC SILICONE |  |  | CFI60Objective:0:1:P | DIC25XII:0:1:H
MRD73400 | CFI PLAN APO LAMBDA S 40XC SILICONE |  |  | CFI60Objective:0:1:P | DIC40XI:0:1:H
MRD73600 | CFI PLAN APO LAMBDA S 60XC SILICONE |  |  | CFI60Objective:0:1:P | DIC60XIV:0:1:H
MRD73950 | CFI SR HP PLAN APO LAMBDA S 100X C SIL |  |  | CFI60Objective:0:1:P | DIC100XI¦DIC100XI-R:0:1:H
TIRF Objectives
MRD01691 | CFI APO 60X OIL TIRF NA 1.49 WD 0.12MM |  |  | CFI60Objective:0:1:P | DIC60XI¦DIC60XI-R:0:1:H
MRD01991 | CFI APO 100X OIL TIRF NA 1.49 WD 0.12MM |  |  | CFI60Objective:0:1:P | DIC100XI¦DIC100XI-R:0:1:H
Super Resolution Objectives
MRY10060 | CFI SR PLAN APO IR 60X WI NA 1.27 |  |  | CFI60Objective:0:1:P | DIC60XIV¦DIC60XIV-R:0:1:H
MRD01995 | CFI SR HP APO TIRF 100XC NA 1.49 OIL W |  |  | CFI60Objective:0:1:P | DIC100XI¦DIC100XI-R:0:1:H
Super Resolution Objectives - Auto Correction Collar
MRD07654 | CFI SR PLAN APO IR AC 60XWI |  | Requires MEP59396 Nosepiece | CFI60ACCObjective:1:1:P
MRD01994 | CFI SR APO TIRF AC 100XH |  | Requires MEP59396 Nosepiece | CFI60ACCObjective:1:1:P
MRY10059 | CFI SR APO 100X OIL TIRF NA 1.49 |  | Older version - Limited quantiy of demo stock only | CFI60Objective:0:1:P
MRD01996 | CFI HP APO TIRF AC 100X OIL |  | Requires MEP59396 Nosepiece | CFI60ACCObjective:1:1:P
Water Dipping Objectives
MRH07120 | CFI PLAN FLUOR 10X W NA 0.3 WD 3.5MM |  |  | CFI60Objective:0:1:P | DIC10X:0:1:H
MRF07420 | CFI FLUOR 40X W NA 0.80 WD 2.0MM |  |  | CFI60Objective:0:1:P | DIC40XIII:0:1:H
MRD07420 | CFI APO 40XW NIR |  |  | CFI60Objective:0:1:P | DIC40XIII:0:1:H
MRD07620 | CFI APO 60XW NIR |  |  | CFI60Objective:0:1:P | DIC60XI:0:1:H
MRF07620 | CFI FLUOR 60X W NA 1.00 WD 2.0MM |  |  | CFI60Objective:0:1:P | DIC60XI:0:1:H
MRL07920 | CFI PLAN 100X W NA 1.1 WD |  |  | CFI60Objective:0:1:P | DIC100XIII:0:1:H
MRP07220 | CFI75 LWD 16X W NA 0.8 WD 3.0 |  |  | CFI75Objective:0:1:P | DIC16XI:0:1:H
MRD77226 | CFI75 APO LWD 20X C W NA 1.00, WD 2.8MM |  |  | CFI75Objective:0:1:P
MRD77220 | CFI75 APO 25X W MP |  |  | CFI75Objective:0:1:P | DIC25XI:0:1:H
MRD77225 | CFI75 APO 25X W MP IR 1300NM |  |  | CFI75Objective:0:1:P | DIC25XI:0:1:H
Specialty and No Cover Glass Objectives
MRD71120 | CFI PLAN APO LWD 10X G, NA 0.5, WD 5.5MM |  |  | CFI60Objective:0:1:P
MRD71220 | CFI90 APO 20X C GLYC NA 1.00, WD 8.2MM |  |  | CFI90Objective:0:1:P
MRL03401 | CFI PLAN ACHRO NCG 40X NA 0.65 WD 0.48MM |  |  | CFI60Objective:0:1:P | DIC40XI¦DIC40XI-C:0:1:H
MRD07650 | CFI PLAN APO IR 60X WI 1.27NA 0.17MM WD |  |  | CFI60Objective:0:1:P | DIC60XIV¦DIC60XIV-R:0:1:H
MRL03901 | CFI PLAN ACHRO NCG 100X NA0.9 WD 1.0MM |  |  | CFI60Objective:0:1:P
MRD04901 | CFI PLAN APO 100X NA 1.4 Oil NCG 0.13mm WD |  |  | CFI60Objective:0:1:P | DIC100XII¦DIC100XII-R:0:1:H
Dispersion Staining Objectives
MRP99101 | CFI ACHROMAT 10X PH DISPERSION STAIN-NC |  |  | CFI60Objective:0:1:P
MRP99200 | CFI PLAN ACHRO 10X NA 0.25 WD 13MM C-DS |  |  | CFI60Objective:0:1:P
MRH99401 | CFI PLAN FLUOR 40X DS DISPERSION STAIN |  |  | CFI60Objective:0:1:P
Strain Free Polarizing Objectives
MRP50042 | CFI P ACHRO 4X-NC |  |  | CFI60Objective:0:1:P
MRP50102 | CFI P ACHRO 10X-NC |  |  | CFI60Objective:0:1:P
MRP50202 | CFI P ACHRO 20X-NC |  |  | CFI60Objective:0:1:P
MRP50402 | CFI P ACHRO 40X-NC |  |  | CFI60Objective:0:1:P
MRP51902 | CFI P ACHRO 100X OIL-NC |  |  | CFI60Objective:0:1:P
Nikon Adv. Modulation Contrast (NAMC)
MRP60105 | CFI Achro NAMC 10X NA |  |  | CFI60Objective:0:1:P | NAMC10x:1:1:H
MRP66205 | CFI ACHROMAT LWD NAMC 20X F-NC |  |  | CFI60Objective:0:1:P | NAMC20x:1:1:H
MRP65425 | CFI ACHROMAT LWD NAMC 40X C-NC |  |  | CFI60Objective:0:1:P | NAMC40x:1:1:H
MRH68200 | CFI S PLAN FLUOR ELWD NAMC 20X C |  |  | CFI60Objective:0:1:P | NAMC20x:1:1:H
MRH68400 | CFI S PLAN FLUOR ELWD NAMC 40X C |  |  | CFI60Objective:0:1:P | NAMC40x:1:1:H
Miscellaneous
MXA20750 | C-OA 15MM OBJECTIVE ADAPTER
MBW10020 | OBJECT MARKER (PURPLE) |  | MXA20750 C-OA 15mm Objective Adapter required for use on all Eclipse CFI60 microscopes
MBW10000 | OBJECTIVE MARKER (RED) |  | MXA20750 C-OA 15mm Objective Adapter required for use on all Eclipse CFI60 microscopes
MXA22018 | CFI SECOND LENS UNIT TUBE LENS ELEMENT
77033028 | OBJECTIVE ADAPTER, RMS M TO CFI F
DIC objective prisms
MBH76210 | D-C DIC SLIDER 10X | DIC |  | DIC10X:1:1:P | DICPrismN1:1:1:H
MBH76220 | D-C DIC SLIDER 20X | DIC |  | DIC20X:1:1:P | DICPrismN2:1:1:H
MBH76221 | D-C DIC SLIDER 20X I-C | DIC |  | DIC20XI-C:1:1:P | DICPrismN1:1:1:H
MBH76521 | T-C DIC SLIDER 20X II | DIC |  | DIC20XII:1:1:P | DICPrismN1:1:1:H
MBH76222 | D-C DIC SLIDER 20X III | DIC |  | DIC20XIII:1:1:P | DICPrismN2:1:1:H
MBH76223 | D-C DIC SLIDER 20X IV | DIC |  | DIC20XIV:1:1:P | DICPrismN2:1:1:H
MBH76230 | D-C DIC SLIDER 25X II | DIC |  | DIC25XII:1:1:P | DICPrismN2:1:1:H
MBH76240 | D-C DIC SLIDER 40X I | DIC |  | DIC40XI:1:1:P | DICPrismN2:1:1:H
MBH76241 | D-C DIC SLIDER 40X I-C | DIC |  | DIC40XI-C:1:1:P | DICPrismN1:1:1:H
MBH76245 | D-C DIC SLIDER 40X II | DIC |  | DIC40XII:1:1:P | DICPrismN2:1:1:H
MBH76640 | D-C DIC SLIDER 40X III | DIC |  | DIC40XIII:1:1:P | DICPrismN2:1:1:H
MBH76540 | T-C DIC SLIDER 40X IV | DIC |  | DIC40XIV:1:1:P | DICPrismN1:1:1:H
MBH76160 | D-C DIC SLIDER 60X I | DIC |  | DIC60XI:1:1:P | DICPrismN2:1:1:H
MBH76162 | D-C DIC SLIDER 60X I-R | DIC |  | DIC60XI-R:1:1:P | DICPrismNR:1:1:H
MBH76260 | D-C DIC SLIDER 60X II | DIC |  | DIC60XII:1:1:P | DICPrismN2:1:1:H
MBH76560 | T-C DIC SLIDER 60X III | DIC |  | DIC60XIII:1:1:P | DICPrismN1:1:1:H
MBH76262 | D-C DIC SLIDER 60X II-R | DIC |  | DIC60XII-R:1:1:P | DICPrismNR:1:1:H
MBH76264 | D-C DIC SLIDER 60X IV | DIC |  | DIC60XIV:1:1:P | DICPrismN2:1:1:H
MBH76266 | D-C DIC SLIDER 60X IV-R | DIC |  | DIC60XIV-R:1:1:P | DICPrismNR:1:1:H
MBH76190 | D-C DIC SLIDER 100X I | DIC |  | DIC100XI:1:1:P | DICPrismN2:1:1:H
MBH76192 | D-C DIC SLIDER 100XI-R | DIC |  | DIC100XI-R:1:1:P | DICPrismNR:1:1:H
MBH76290 | D-C DIC SLIDER 100X II | DIC |  | DIC100XII:1:1:P | DICPrismN2:1:1:H
MBH76292 | D-C DIC SLIDER 100XII-R | DIC |  | DIC100XII-R:1:1:P | DICPrismNR:1:1:H
MBH76490 | D-C DIC SLIDER 100X III | DIC |  | DIC100XIII:1:1:P
DIC objective prisms for MBP70110 Single Position Nosepiece
MBH76720 | D-C SLIDER 16X I | DIC |  | DIC16XI:1:1:P | DICPrismN2:1:1:H
MBH76225 | D-C DIC SLIDER 25X I | DIC |  | DIC25XI:1:1:P | DICPrismN2:1:1:H
NAMC Modules
MEH51580 | TC-C-MN-10X NAMC 10X MODULE |  |  | CondenserModule:0:1:P | NAMC10x:1:6:P
MEH51600 | TC-C-MN-20X NAMC 20X MODULE |  |  | CondenserModule:0:1:P | NAMC20x:1:6:P
MEH51620 | TC-C-MN-40X NAMC 40X MODULE |  |  | CondenserModule:0:1:P | NAMC40x:1:6:P
MEH51590 | TC-C-MNL-10X NAMC LWD 10X MODULE |  |  | CondenserModule:0:1:P | NAMC10x:1:6:P
MEH51610 | TC-C-MNL-20X NAMC LWD 20X MODULE |  |  | CondenserModule:0:1:P | NAMC20x:1:6:P
MEH51630 | TC-C-MNL-40X NAMC LWD 40X MODULE |  |  | CondenserModule:0:1:P | NAMC40x:1:6:P
TU Objectives
MUC11150 | TU PLAN APO EPI 150X |  |  | CFI60Objective:0:1:P
MUE12030 | T PLAN EPI 2.5X |  |  | CFI60Objective:0:1:P
MUE12050 | TU PLAN FLUOR EPI 5X |  |  | CFI60Objective:0:1:P
MUE12100 | TU PLAN FLUOR EPI 10X |  |  | CFI60Objective:0:1:P
MUE12200 | TU PLAN FLUOR EPI 20X |  |  | CFI60Objective:0:1:P
MUE12500 | TU PLAN FLUOR EPI 50X |  |  | CFI60Objective:0:1:P
MUE12900 | TU PLAN FLUOR EPI 100X |  |  | CFI60Objective:0:1:P
MUE13050 | TU PLAN FLUOR EPI P 5X |  |  | CFI60Objective:0:1:P
MUE13100 | TU PLAN FLUOR EPI P 10X |  |  | CFI60Objective:0:1:P
MUE13200 | TU PLAN FLUOR EPI P 20X |  |  | CFI60Objective:0:1:P
MUE13500 | TU PLAN FLUOR EPI P 50X |  |  | CFI60Objective:0:1:P
MUE13900 | TU PLAN FLUOR EPI P 100X |  |  | CFI60Objective:0:1:P
MUE21200 | TU PLAN FLUOR EPI ELWD 20X |  |  | CFI60Objective:0:1:P
MUE21500 | TU PLAN FLUOR EPI ELWD 50X |  |  | CFI60Objective:0:1:P
MUE31500 | T PLAN FLUOR EPI SLWD 50X |  |  | CFI60Objective:0:1:P
MUE42050 | TU PLAN FLUOR BD 5X |  |  | CFI60Objective:0:1:P
MUE42200 | TU PLAN FLUOR BD 20X |  |  | CFI60Objective:0:1:P
MUE42900 | TU PLAN FLUOR BD 100X |  |  | CFI60Objective:0:1:P
MUE61200 | TU PLAN BD ELWD 20X |  |  | CFI60Objective:0:1:P
MUE61500 | TU PLAN BD ELWD 50X |  |  | CFI60Objective:0:1:P
MUE61900 | TU PLAN BD ELWD 100X |  |  | CFI60Objective:0:1:P

## NIS Elements
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
Core Packages
MQS34000 | NIS-ELEMENTS HIGH CONTENT HC PACKAGE |  | * Multiwell plate acquisition, databasing, analysis and review package, includes camera and Nikon motorized microscope drivers (except Nikon RFA). 6D Imaging, Live/ Dead, Cell Counting, General Analysis and JOBS. *Core is based on AR package and requires AR SUA (MQS31410) Note: Separate drivers are required for 3rd party device control. Note: It is possible to upgrade an exising Ar or C package to HC using MQS43050, but 6D must be quoted or included on original license. | NISPackage:0:99:H
MQS31000 | NIS-ELEMENTS AR |  | * Included modules: Advanced Interpreter, Live Compare, and Time Measurements, Automeasure, HDR * Optional modules include: 6D acquisition, JOBS, Illumination Sequence, General Analysis, Real time Deconvolution, LIM Decon Tools, EDF, 2D Object Tracking, 3D Measurement, Ratio & FRET, Object Classifier Note: Does not include 4D or 6D capabilities, must quote 6D for any multidimensional imaging Note: Separate drivers are required for 3rd party device control. Note: It is possible to upgrade a D or a BR package to an AR. | NISPackage:0:99:H
MHS50000 | NIS-ELEMENTS CONFOCAL PACKAGE |  | NIS-Elements C Imaging Software for AX, A1 and C2 Confocal Applications. Includes 6D Imaging Module (MQS42560). SUA for Confocal Package is only available through Service Contract or Preventative Maintenance (PM) through Service Dept. Separate drivers are required for 3rd party device control. CANNOT upgrade an AR or HC(AR) package to the Confocal Package. | NISPackage:0:99:H
MQS32000 | NIS-ELEMENTS BR |  | * Optional modules include Advanced Interpreter, Live Compare, Time Measurement, Image Database, EDF, 4D Acquisition, HDR Note: It is possible to upgrade a D to a BR package. | NISPackage:0:99:H
MQS33000 | NIS-ELEMENTS D |  | Single document interface. * Optional Modules include: Automeasure, Advanced Interpreter, Live Compare, Image Database, EDF, HDR and Metalo Modules | NISPackage:0:99:H
77010629 | NIS-ELEMENTS C NETWORK LICENSE |  |  | NISPackage:0:99:H
77010630 | NIS-ELEMENTS HC NETWORK LICENSE |  |  | NISPackage:0:99:H
77010631 | NIS-ELEMENTS AR ACTIVE NETWORK LICENSE |  |  | NISPackage:0:99:H
Multi License (ML) Packages
MQS31100 | NIS-ELEMENTS AR MULTI-LICENSE |  | Additional license for Acquisition and analysis package. * Optional modules include: 6D acquisition, Image Database, Decon Tools, EDF, 2D Object Tracking, Ratio & FRET, Object Classifier, General Analysis. Note: Does not include 4D capabilities, must quote 6D for any multidimensional imaging Note: Separate drivers are required for 3rd party device control. Note: The multi-license can only be sold to an end user who has a registered copy of AR (Material # MQS31000) or C (Material # MHS50000). | NISPackage:0:99:H
MQS32100 | NIS-ELEMENTS BR MULTI-LICENSE |  | Additional license for Acquisition and analysis package. * Optional modules include: Advanced Interpreter, Live Compare, Time Measurement, Image Database, 4D Acquisition, & HDR Note: Separate drivers are required for 3rd party device control. Note: The multi-license can only be sold to an end user who has a registered copy of BR (Material # MQS32000). End user cannot purchase BR ML if they own only AR. | NISPackage:0:99:H
MQS33100 | NIS-ELEMENTS DOCUMENTATION MULTI LICENSE |  | Additional license for Acquisition and analysis package. Single document interface. * Optional Modules include: Automeasure, Advanced Interpreter, Live Compare, Image Database, EDF, HDR and Metalo Modules. * Includes one of the following: Time Lapse, Multipoint Acquisition, or Z control. Note: Additional drivers are required for 3rd party hardware device control. Note: The multi-license can only be sold to an end user who has a registered copy of D (Material # MQS33000). | NISPackage:0:99:H
77010632 | NIS-ELEMENTS AR ML NETWORK LICENSE |  |  | NISPackage:0:99:H
Passive / Offline Package
MQS31500 | NIS-ELEMENTS AR STAND ALONE PASSIVE |  | Analysis, processing and visualization only, no device control available. Note: Additional software modules available for purchase. Note: It is possible to sell AR Passive to a customer without an exisiting license. Note: It is not possible to upgrade an AR Passive to an AR Active Key | NISPackage:0:1:P
77010609 | NIS-ELEMENTS AR PASSIVE BUNDLE |  | Includes AR Passive, 2D/3D Tracking and GA3. | NISPackage:0:1:P
77010611 | NIS-ELEMENTS NETWORK HASP |  |  | NISPackage:0:1:P
77010626 | NIS-ELEMENTS AR PASSIVE NETWORK 1 LIC |  |  | NISPackage:0:1:P
77010627 | NIS-ELEMENTS AR PASSIVE NETWORK 3 LIC |  |  | NISPackage:0:1:P
77010628 | NIS-ELEMENTS AR PASSIVE NETWORK 5 LIC |  |  | NISPackage:0:1:P
77010641 | NIS-ELEMENTS NETWORK SEAT FOR AR PASSIVE |  | Adds additional seats to an AR passive network license. For existing network key licenses, they must be current with their SUAs to add additional licenses. Do NOT use MQS31500 to add additional licenses to network keys. | NISPackage:0:1:P
AR DUO Package
MQS31200 | NIS-ELEMENTS AR DUO |  | * 1 License of Advanced Research (Material # MQS31000) bundled with 1 Passive Advanced Research license (Material # MQS31500). * 2 separate keys will be issued, AR SUA (MQS31410) needed for both keys | NISPackage:0:99:H
Software Modules
MQS42100 | ELEMENTS: EXTENDED DEPTH FOCUS |  | * Combines multiple focal images into single focused 2D image. * Surface View, Z Profile, 3D Length Measurements included with module. | NISPackage:0:1:P
MQS42300 | NIS-AMEAS AUTO MEASURE MODULE |  |  | NISPackage:0:1:P
MQS42690 | NIS-ELEMENTS: LIM 2D DECON TOOLS |  | * LIM's 2D Decon Suite: 2D Decon, 2D Fast (Wiener Filter for Live or Post Processing) Required for Advanced and Live Denoising in Version 4.5. Available for C, HC & AR only. Includes Clarify. ai for Version 5.3 and higher. | NISPackage:0:1:P
MQS42700 | NIS-ELEMENTS: LIM 2D/3D DECON TOOLS | NIS-3D-DECON | *LIM's 2D & 3D Decon Suite: 2D Decon, 2D Fast Decon (Wiener Filter for Live or Post Processing), 3D Blind, 3D PSF, 3D Wiener Filter Enables Advanced and Live Denoising in Version 4.5 (Includes part # MQS42690) Available for C, HC & AR only. Includes Clarify. ai for Version 5.3 and higher. | NISPackage:0:1:P
MQS42790 | NIS-ELEMENTS: HDR |  | High Dynamic Range for using multiple exposures to optimize range. Included in C, HC, Advanced Research, optional module for Basic Research and Documentation. | NISPackage:0:1:P
MQS42410 | ELEMENTS: IMAGE DATABASE |  | * Uses MS Access engine, Optional module for C, AR, BR and D. | NISPackage:0:1:P
MQS42510 | ELEMENTS: ADVANCED INTERPRETER API |  | * Advanced macro functionality. C like command language. * Available for Basic Research and Documentation packages. Included in Advanced Research. | NISPackage:0:1:P
MQS42520 | ELEMENTS: LIVE COMPARE |  | * Real time image comparator with saved images. * Available for Basic Research and Documentation packages. Included in C, HC & AR. | NISPackage:0:1:P
MQS42530 | ELEMENTS: TIME MEASUREMENT |  | * Adds time analysis capabilities to Basic Research package. Note: Does not include Ratio View or functionality, 4D module does not include Time Measurements Available for Basic Research only. Included in C, HC, and AR packages. | NISPackage:0:1:P
MQS42570 | NIS-ELEMENTS MODULE: OBJECT TRACKING |  | * Image analysis module: 2D object tracking of distance, acceleration, velocity and path length. Various overlay/annotation options for tracking included. * Includes: time, channel or wavelength, focus, multipoint, or multi stage position. Available for C, HC & AR only. | NISPackage:0:1:P
MQS42580 | NIS-ELEMENTS MODULE: RATIO & FRET |  | * Ratio: Ratio View, Live Ratio Graphing & Measurement, Calcium & New Titration Calibration * FRET: capture and analysis Available for C, HC & AR only. | NISPackage:0:1:P
MQS42590 | NIS-ELEMENTS MODULE: OBJECT CLASSIFIER |  | * Advanced analysis module for classifying objects based on shape, intensity, area, etc Available for C, HC & AR only. | NISPackage:0:1:P
MQS42610 | NIS-ELEMENTS METALO MODULE: GRAIN SIZING |  | Image analysis module for measuring steel grain size as per ASTM Specification Note: NIS-Elements D (MQS33000) & Automeasure (MQS42300) are required. | NISPackage:0:1:P
MQS42620 | NIS-ELEMENTS METALO MODULE: CAST IRON |  | Image analysis module for detection and classification of standard cast iron nodule types Note: NIS-Elements D (MQS33000) & Automeasure (MQS42300) are required. | NISPackage:0:1:P
MQS42900 | NIS-ELEMENTS 2D ADVANCED TRACKING |  | Optimized for large number of objects, splitting objects, tracking using binary objects, probability ellipse, vector graph Available for C, HC & AR only. | NISPackage:0:1:P
MQS42920 | NIS-ELEMENTS MODULE: FILTER ANALYSIS |  | Industrial analysis module for particles trapped on standardized & custom size fluid filters. Automated scanning & image capture. Comprehensive reporting. Supports ISO 16232. Available for all core packages. MQS41100, MQS42300 (for D) required. | NISPackage:0:1:P
MQS42940 | NIS-ELEMENTS MODULE: 3D MEASUREMENT |  | 3D thresholding, 3D viewing and voxel based measurements, centroid calculation of binary Available for C, HC & AR only. | NISPackage:0:1:P
MQS42950 | NIS-ELEMENTS 2D+3D MEASUREMENT TRACKING |  | Advanced 2D tracking, 3D Thresholding, 3D Viewing, voxel-based measurements and centroid calculation of binary objects. Includes MQS42570, MQS42900, MQS42940 (formerly 77010582) Available for C, HC & AR only. | NISPackage:0:1:P
MQS43130 | NIS-ELEMENTS MODULE: JOBS ACQUISITION |  | Acquistion Module for custom routines. Available for C & Advanced Research (Ar) only. MQS42560 6D Imaging Module required Included in High Content (HC) Package - MQS34000 | NISPackage:0:1:P
MQS43150 | NIS-ELEMENTS MOD: GEN ANALYSIS UPGRADE |  | If a customer already has GA2, this is the upgrade material to GA3. | NISPackage:0:1:P
77010599 | NIS-ELEMENTS MOD: SNGL PARTICLE TRACKING |  | * Image analysis module for quantifying fluorescent molecule dynamics. Single Particle Tracking (SPT) module determines movement and diffusion characteristics from the tracks of individual molecules. Requires MQS42900. Can be analyzed offline, but acquired data must have been generated from an NIS-Elements Illumination Sequence-Triggered system. Please note: CHECK WITH SW DEPT FOR SUPPORTED CAMERAS & DEVICES (i.e. Photometrics Cameras not supported) Available for C, HC & AR only. | NISPackage:0:1:P
77010604 | NIS-ELEMENTS MOD: GEN ANALYSIS 2 AND 3 |  | Image processing and measurement analysis engine for custom and standard routines with support of ND dimensions, 3D features, Javascript support, output controls and many data management functions. This module includes both originally released General Analysis (GA2) (MQS43110) and GA3. (Includes MQS43150 and MQS43110) Available for C, HC,AR or AR Passive Only. | NISPackage:0:1:P
77010607 | NIS-ELEMENTS MOD: MEASUREMENT SEQUENCER |  | A set of measurement tools for the creation of simple to complex features which can then be repeatedly measured by the user. Applicable to both Bio and Industrial applications such as Weld Geometry. Customizable template based workflow for product specific standards. Comprehensive reporting and image annotation. Available for C, AR, BR and D Core packages. | NISPackage:0:1:P
77010612 | NIS-ELEMENTS: BACK APERTURE OBSERVATION | NIS-BACKAPCAM | Module for Ti2 inverted microscope for back aperture observation. Runs in a separate window from the main imaging window. C, AR or HC package required required. Not available for BR or D. Runs in NIS-Elements 5.1 or higher. | NISPackage:0:1:P
77010613 | NIS-ELEMENTS: OPTOGENETICS ACQUISITION |  | Module for behavioral imaging. Will be compatible with A1 detector or any PCO supported camera. C, AR or HC package required. For real time analysis, GA3 77010604 is recommended. | NISPackage:0:1:P
MQS43200 | NIS-ELEMENTS MODULE: NIS.AI MODULE |  | NIS-Elements module for artificial intelligence and deep learning using convolutional neutral networks (CNNs) to learn from training data to apply repeatedly to volumes of data. Includes Convert.ai, Enhance.ai and Segment.ai. Available for HC, C, AR, and AR Passive. Work as part of pipeline with Jobs and GA3, which are available separately. Does not include Clarify.ai. | NISPackage:0:1:P
Offline Deconvolution and Batch Denoise Options
77010575 | NIS-ELEMENTS LIM2D/3D BATCHDECON OFFLINE |  | Separate offline package with a specific HASP key for scheduling multiple deconvolution tasks. Offline Decon operates at NIS Version 6.10. It takes advantage of GPU processing and latest Volume Viewer options. | NISPackage:0:1:P
77010576 | NIS-ELEMENTS 2 COPIES BATCH DECON OFFLIN |  | Separate offline package with a specific HASP key for scheduling multiple deconvolution tasks. Offline Decon operates at NIS Version 6.10. It takes advantage of GPU processing and latest Volume Viewer options. Ships with 2 HASP keys. | NISPackage:0:1:P
77010577 | NIS-ELEMENTS 3 COPIES BATCH DECON OFFLIN |  | Separate offline package with a specific HASP key for scheduling multiple deconvolution tasks. Offline Decon operates at NIS Version 6.10. It takes advantage of GPU processing and latest Volume Viewer options. Ships with 3 HASP keys. | NISPackage:0:1:P
77010578 | NIS-ELEMENTS 5 COPIES BATCH DECON OFFLIN |  | Separate offline package with a specific HASP key for scheduling multiple deconvolution tasks. Offline Decon operates at NIS Version 6.10. It takes advantage of GPU processing and latest Volume Viewer options. Ships with 5 HASP keys. | NISPackage:0:1:P
77010579 | NIS-ELEMENTS ADD BATCH OFFLINE LIM DECON |  | Adds batching interface to NIS-Elements C, HC, AR Package with Existing MQS42700 2D/ 3D Decon module Offline Decon operates at NIS Version 6.10. It takes advantage of GPU processing and latest Volume Viewer options. | NISPackage:0:1:P
77010620 | NIS-ELEMENTS BATCH DENOISE LICENSE |  | Separate offline package with a specific HASP key for scheduling multiple denoising tasks. Offline Denoise operates at NIS Version 6.02.01 (build 1955). | NISPackage:0:1:P
77010621 | NIS-ELEMENTS BATCH DENOISE 2 LICENCES |  | Separate offline package with a specific HASP key for scheduling multiple denoising tasks. Offline Denoise operates at NIS Version 6.02.01 (build 1955). Ships with 2 HASP keys. | NISPackage:0:1:P
77010622 | NIS-ELEMENTS BATCH DENOISE 3 LICENCES |  | Separate offline package with a specific HASP key for scheduling multiple denoising tasks. Offline Denoise operates at NIS Version 6.02.01 (build 1955). Ships with 3 HASP keys. | NISPackage:0:1:P
77010623 | NIS-ELEMENTS BATCH DENOISE 5 LICENCES |  | Separate offline package with a specific HASP key for scheduling multiple denoising tasks. Offline Denoise operates at NIS Version 6.02.01 (build 1955). Ships with 5 HASP keys. | NISPackage:0:1:P
Hardware Control Modules
MQS41100 | NIS ELEMENTS DRIVER FOR MOTORIZED STAGES | NIS-3PSTAGES | * Device control for Prior, ASI, Ludl and Maerzhauser stages. Note: Nikon Ti-E Motorized Stage does not require software driver. | NISPackage:0:1:P
MQS41220 | ELEMENTS: SUTTER/UNIBLITZ FILTRWHL+SHUTR | NIS-FILTERWHEEL | * Device control for Sutter, Prior, ASI, Ludl & Uniblitz illumination devices including shutters and filter wheels. Note: Required for Use of General Shutter and General Filters Option Note: Adds SHUTTER & FILTER bits to HASP key | NISPackage:0:1:P
MQS41320 | ELEMENTS:PRIOR/MAERZHAUSER/RFA/PI-PIEZO | NIS-PIEZO | * Device control for external motorized focus control. * Required for Nikon RFA, Prior, ASI, Ludl, Maerzhauser, and PI Piezo Z drive (Note that piezo is via RS232 or through voltage triggering). If triggering, MQS42780 required. | NISPackage:0:1:P
MQS41410 | ELEMENTS:CAM EMISSIONS SPLITTER-DUAL VIE | NIS-SPLITTER | * Software module for splitting and displaying real time images from any emission splitting device such as a Dual View. | NISPackage:0:1:P
MQS41450 | NIS-ELEMENTS HARDWARE MODULE: DUALCAMER | NIS-DUALCAM | * Hardware module for allowing simultaneous camera control. Needed for any Dual Andor, Dual Hamamatsu Flash 4.0. Available for C, HC & AR only. | NISPackage:0:1:P
MQS41600 | NIS-ELEMENTS ROBOTIC WELL PLATE LOADER |  | High Content Robotic Driver for automated delivery of wellplates in the hotel stackers to and from the microscope stage. Available for C, HC & AR only. C & Ar copies must have Jobs (MQS43130). | NISPackage:0:1:P
MQS41610 | NIS-ELEMENTS MODULE FOR LINKAM HEATED ST |  | * Heated Stage Control for Linkam Stage Modules: TP 94, 93, 92 Available for all core packages. | NISPackage:0:1:P
MQS41920 | NIS ELEMENTS: EXFO ILLUMINATOR | NIS-EXFO | * Controls shutter and illumination intensity from EXFO light source and MANY other light sources Notes: Required Driver for Sutter Lambda XL, Lumencor Sola, Lumen Dynamics X-Cite and other devices. Check Parts by Package list for full list of devices requiring this driver. Note: Adds SHUTTER & LIGHT bits to HASP key | NISPackage:0:1:P
MQS41930 | NIS-ELEMENTS: WAVELENGTH SWITCHER | NIS-WLSWITCHER | * Device control for of Sutter DG4/5, AOTF, Agilent, Lumencor SpectraX, Spinning Disk, etc. Check Parts by Package list for full list of devices requiring this driver. * DOES NOT Include the Wavelength Control Module (MQS41220), which provides 'FILTER' to the HASP key, which would be needed for any 3rd party filter wheel Was previously Material # 77010560 (which did include MQS41220) Note: LUSU Control is now a part of the NIS-Elements Core AR Package. Check Parts by Package list for full list of devices requiring this driver. Note: Adds SHUTTER &WAVELENGTH SWITCHER bits to HASP key | NISPackage:0:1:P
MQS41940 | NIS-ELEMENTS: DAQ TTL/ANALOG IO | NIS-TRIGGERING | *Driver for TTL/Voltage Control in ND experiments *Quote MQS42560 (6D) for Ar or , 6D is included in C Not available for Br or D packages Check Parts by Package list for full list of devices requiring this driver. Check Camera and Device List posted with each new release to check for NIDAQ requirements and device selection. | NISPackage:0:1:P
MQS41950 | NIS-ELEMENTS MODULE: YOKOGAWA CSU-X1 | NIS-X1 | * Device control for Yokogawa CSU X1 Check Parts by Package list for full list of modules required for device control. Available for C, HC & AR only. | NISPackage:0:1:P
MQS41960 | NIS-ELEMENTS MODULE: YOKOGAWA CSU-W1 | NIS-W1 | * Device control for Yokogawa CSU W1 Check Parts by Package list for full list of modules required for device control. Available for C, HC & AR only. | NISPackage:0:1:P
77010640 | NIS-ELEMENTS MODULE: CREST CONFOCAL | NIS-CREST | Device control for Crest Confocal Systems (V2, V3 and Cicero). ONLY Compatible with NIS-Elements v6.02.01 or later. Available for C, HC & AR only. | NISPackage:0:1:P
MQS41970 | NIS-ELEMENTS MODULE: XY GALVO | NIS-XYGALVO | The updated material number for the former 77010594. | NISPackage:0:1:P
MQS42540 | ELEMENTS: 4D IMAGING | NIS-4D | * Acquisition module for up to 4 dimensions of image acquisition, including any two of the following: time, multi-stage point, channel/wavelength and Z focus. Note: 3rd party device drivers are required. Note: Module does not include Time Measurements Available for Basic Research only. Note: Module does not have Large Image Stitching in ND (Available in Advanced Research + 6D Only) | NISPackage:0:1:P
MQS42560 | ELEMENTS: 6D IMAGING | NIS-6D | * Acquisition module for 6 dimensional imaging * Includes: time, channel or wavelength, focus, multipoint, or multi stage position. * Includes: ND Sequence, ND Multi Position for Nested Experiments Note: 3rd party device drivers are required. Available for Advanced Research only. Included in C (MHS50000) | NISPackage:0:1:P
MQS42780 | NIS-ELEMENTS: TRIGGERED DEVICE CONTROL | NIS-TRIGGER | * Device control for externally triggering a device through a camera *MQS41320 is a required module for triggered Piezo Z *MQS41930 is a required module for triggered DG4/5 Available for C, HC, AR & BR only. | NISPackage:0:1:P
MQS42930 | NIS-ELEMENTS MODULE: ILLUMINATION SEQUE |  | *Illumination Sequence for flexible triggered laser and LED experiment set up *Requires MQS42560 6D and MQS42780 Triggered Acquisition modules *Please note: CHECK WITH SW DEPT FOR SUPPORTED CAMERAS & DEVICES *Required for STORM with Version 5.0 and higher Available for C, HC & AR only. | NISPackage:0:1:P
77010598 | NIS-ELEMENTS MODULE: MIGHTEX POLYGON DMD | NIS-POLYGON | * Driver for Digital Mirror Device MighTex Polygon 400 & Polygon 1000. Available for C, HC & AR only. | NISPackage:0:1:P
77010601 | NIS-ELEMENTS MOD: ADD ON TRIP/QUAD CAM | NIS-TRIPCAM | * Upgrade ADD-ON module for Triple/ Quad Camera Requires MQS41450- Supported in C, HC, AR Packages only. Supported modules: Andor Ixon 897, X3, 897 Ultra, Ixon 885. Available for C, HC & AR only. | NISPackage:0:1:P
97172 | NIS-ELEMENTS MODULE: SLIDE LOADER |  | * Device Control for Prior PL200 * Note: If ordering slide loader module, only device control for moving a slide on or off the microscope stage. Scan Large Image with Z Dialog only workflow is included. Available for C, HC, AR & BR (MQS42510 needed for BR) | NISPackage:0:1:P
97173 | NIS-ELEMENTS MODULE: INCUBATION CONTROL |  | * Device Control for Microscope Incubation Chamber- Check with NII SW Dept. for compatible devices Available for C, HC, AR & BR (API and 4D needed for Br, 6D needed for Ar.) | NISPackage:0:1:P
77010618 | NIS-ELEMENTS MODULE: FLUOR LIFETIME IMAG |  | NIS-ELEMENTS MODULE- Fluorescence Life Time Imaging for Becker & Hickl for A1 - See A1 MP Product Sheet for related materials and more information. | NISPackage:0:1:P
77010633 | NIS-ELEMENTS iLAS2 TIRF DRIVER | NIS-ILAS |  | NISPackage:0:1:P
77010637 | NIS-ELEMENTS MODULE: DeepSIM | NIS-DEEPSIM | Device Control for CrestOptics DeepSIM. Required for both DeepSIM stand-alone and in combination wiith Crest X-Light V3. | NISPackage:0:1:P
STORM / SIM Modules
MQS42650 | NIS-ELEMENTS: STORM ACQ & ANALYSIS |  | * Acquisition & Analysis module for STORM *Illumination Sequence module needed MQS42930 and triggering MQS42780 Available for AR, AR ML, AR DUO, HC and C packages. | NISPackage:0:1:P
MQS42660 | NIS-ELEMENTS MODULE STORM OFF ANALYSIS |  | *Offline Analysis module for STORM -must be added to a HASP key such as MQS35100 AR Passive or MQS33100 AR ML Requires the purchase of MQS42650 Available for AR, AR ML, AR DUO, AR Passive, HC and C packages. | NISPackage:0:1:P
MQS42670 | NIS-ELEMENTS MODULE: SIM ACQ & ANALYSIS |  | * Acquisition & Analysis module for SIM Available for AR, AR ML, AR DUO, HC and C packages. | NISPackage:0:1:P
MQS42680 | NIS-ELEMENTS MODULE SIM OFF ANALYSIS |  | *Offline Analysis module for SIM -must be added to a HASP key such as MQS35100 AR Passive or MQS33100 AR ML Requires the purchase of MQS42670 Available for AR, AR ML, AR DUO, AR Passive, HC and C packages. | NISPackage:0:1:P
NIS-Elements Documentation Imaging Source Bundle
97164 | NIS-ELEMENTS D W/ BASLER BUNDLE |  | *NIS-Elements Documentation + Imaging Source USB 3.0 1920 x 1200 full HD Camera + camera cable Camera only available in US with purchase of NIS-Elements Documentation (included in Bundle) Imaging Source Camera will only work in the NIS-Elements Documentation (D) Package. Camera not supported in C, AR, BR or HC. If this D core package is upgraded to BR, the camera will not be supported. https://na27.salesforce.com/50133000000ixcw
Package and Module Upgrades / Reconfigfuration
MQS31210 | NIS-ELEMENTS DOCUMENTATION TO AR UPGRADE |  | Note: When this upgrade is applied, the SUA date will be reset and the key will now require AR SUAs. | NISPackage:0:1:P
MQS32210 | NIS-ELEMENTS DOCUMENTATION TO BR UPGRADE |  | Note: When this upgrade is applied, the SUA date will be reset to current date and the key will now require BR SUAs. | NISPackage:0:1:P
MQS31220 | ELEMENTS SOFTWARE UPGRADE BR TO AR |  | When this upgrade is applied, the SUA date will be reset to current date and the key will now require AR SUAs. If the customer's key already has the 4D module, MQS31230 is required. | NISPackage:0:1:P
MQS31230 | NIS-ELEMENTS UPGRADE: 4D TO 6D IMAGING |  | * Please note that this material is REQUIRED if the existing BR copy that will be upgraded already has the 4D imaging module. | NISPackage:0:1:P
MQS43050 | NIS-ELEMENTS HC UPGRADE FOR AR & C w/GA3 |  | The updated HC Upgrade material for the former MQS43040. MQS43050 includes GA3. Does not reset SUA date. | NISPackage:0:1:P
77010591 | NIS-ELEMENTS HASP RECONFIGURATION FEE |  | This fee is for Module transfers between two HASP keys that are at the same version access (same SUA level). HASP key module transfers are not possible between keys with different version access. Always check SUA dates of both keys using the HASP Checker Site before quoting this material #. Module transfers are allowed between different types of keys (C, AR, AR Passive, BR, D) as long as the given module is compatible with the core license that will be trans | NISPackage:0:1:P
MQS44000 | NIS-ELEMENTS: SMART EXPERIMENT ASSAYS |  |  | NISPackage:0:1:P
MQS44010 | NIS-ELEMENTS NUCLEAR TRANSLOCATION ASSAY |  |  | NISPackage:0:1:P
MQS44040 | NIS-ELEMENTS DNA DAMAGE GAMMA-H2AX ASSAY |  |  | NISPackage:0:1:P
MQS44020 | NIS-ELEMENTS APOPTOSIS ASSAY |  | Supported objective for this Assay: CFI60 Plan Apochromat Lambda D 10x (MRD70170) | NISPackage:0:1:P
MQS44030 | NIS-ELEMENTS AUTOPHAGY ASSAY |  | Supported objective for this Assay: CFI60 Plan Apochromat Lambda D 20x (MRD70270) | NISPackage:0:1:P
MQS44050 | NIS-ELEMENTS MICRONUCLEUS TEST ASSAY |  | Supported objective for this Assay: CFI60 Plan Apochromat Lambda D 20x (MRD70270) | NISPackage:0:1:P
MQS44060 | NIS-ELEMENTS ENDOCYTOSIS ASSAY |  | Supported objective for this Assay: CFI60 Plan Apochromat Lambda D 20x (MRD70270) | NISPackage:0:1:P
MQS44070 | NIS-ELEMENTS PHAGOCYTOSIS ASSAY |  | Supported objective for this Assay: CFI60 Plan Apochromat Lambda D 20x (MRD70270) | NISPackage:0:1:P
MQS44080 | NIS-ELEMENTS MITOCHONDRIA TOXICITY ASSAY |  | Supported objective for this Assay: CFI60 Plan Apochromat Lambda D 20x (MRD70270) | NISPackage:0:1:P
NISPackage:0:1:P

## Promos
Product Code | Description | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
999-343 | Ti2A Intelligent Inverted Microscope Promo Package |  | AddPart:MEA54100,MEB55835,MEB52340,MAK10110,MAK10110,MEP53325,MEE59925,MEE55705,MEL51010,MEL56200,MEV51020,MXA22147,MEE54860,MEE54741,MQF52057,97050,79035,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084
999-344 | Ti2E Motorized Microscope Promo Package |  | AddPart:MEA54000,MEF55037,MEB55835,MEB52340,MAK10110,MAK10110,MEP59320,MEE59925,MEE55705,MEL51010,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,97050,79035,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084
999-345 | Ti2E Motorized PFS Microscope Promo Package |  | AddPart:MEA54000,MEF55037,MEB55835,MEB52520,MAK10110,MAK10110,MEP59394,MEE59925,MEE55705,MEL51920,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,97050,79035,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084
999-346 | Ti2E Motorized Stage Promo Package |  | AddPart:MEC56130,MEF55706,MEC59140,MEC59150
999-347 | Ti2E Motorized Encoded Stage Promo Package |  | AddPart:MEC56120,MEF55706,MEC59140,MEC59150
999-348 | Ti2E Fluorescence Promo Package |  | AddPart:MBF83000,MQF52057,79035,97050,96378
999-349 | Ti2E Research Plan Fluor Objective Promo Package |  | AddPart:MRH20045,MRH10105,MRH48230,MRH48430,MRH02902
999-350 | Ti2E Research Plan Apochromat Objective Promo Package |  | AddPart:MRD70040,MRD70170,MRD70270,MRD70470,MRD07602
999-351 | Ti2E Motorized MP PFS Microscope Promo Package |  | AddPart:MEA54000,MEF55037,MEB55835,MEB52520,MAK10110,MAK10110,MEP59395,MEE59925,MEE55705,MEL51920,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,97050,79035,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084
999-352 | JI Assay Promotion Package |  | AddPart:MFA92000,MFD92000,MFF92000,MFC92000,MFD92020,MBF83200,MQS44000,MRD70040,MRD70170,MRD70270,MQF52057,79035,79035,97050,97050,96378,MQS42560,MQS31000,MQS43130,77010604
999-353 | NiE Motorized Research Microscope Promo Package |  | AddPart:MBA91000,MBD92100,MBD93010,MBF95420,MBF95450,MBE65800,79035
999-354 | Motorized Research Plan Fluor Add On Promo Package |  | AddPart:MAK30105,MAK30105,MBL99800,MRH00105,MRH00205,MRH00405,MRH01902
999-355 | Motorized Research Plan Apo Add On Promo Package |  | AddPart:MAK30105,MAK30105,MBL99800,MRD70040,MRD70170,MRD70470,MRD71970
999-356 | Ni Epi-Fl Motorized Imaging Head with Shutter Promo Package |  | AddPart:MBB94511,MBB96710,MBE94150,MBE94300,MBE96000,MBE96910,MBE96930
999-357 | Ni Motorized Epi-Fl Trinoc Head with Shutter Promo Package |  | AddPart:MBB93115,MBE94150,MBE94300,MBE96000,MBE96910,MBE96930
999-358 | NiL Basic Clinical Microscope Promo Package |  | AddPart:79035,MBA92020,MBC91146,MBD92000,MBD93000,MBD94000
999-359 | Complete Plan Fluor Clinical Add On Promo Package |  | AddPart:MAK10110,MAK10110,MBB93800,MBL71305,MBP71316,MRH00045,MRH00105,MRH00405,MXA22048
999-360 | Complete Plan Apo Clinical Add On Promo Package |  | AddPart:MAK10110,MAK10110,MBB93800,MBL71305,MBP71316,MRD70040,MRD70170,MRD70470,MXA22048
999-361 | NiL Basic Research Microscope Promo Package |  | AddPart:79035,MBA92020,MBC91036,MBD92000,MBD93010,MBD94000
999-362 | Manual Research Plan Fluor Add On Promo Package |  | AddPart:MAK30105,MAK30105,MBL99005,MRH00105,MRH00205,MRH00405,MRH01902
999-363 | Manaul Research Plan Apo Add On Promo Package |  | AddPart:MAK30105,MAK30105,MBL99005,MRD70040,MRD70170,MRD70470,MRD71970
999-364 | Ni Manual Epi-Fl - Trinoc Head Promo Package |  | AddPart:MBB93115,MBE92020,MBE94150,MBE94500
999-365 | Ni Epi-Fl Manual Imaging Head Promo Package |  | AddPart:MBB94522,MBB96721,MBE92020,MBE94150,MBE94500
999-366 | NiL Motorization Ready Research Microscope Promo Package |  | AddPart:79035,79035,MBA92020,MBC91035,MBD92110,MBD93010,MBD94000,MBF95410
999-367 | Ts2 Tissue Culture DL Phase Conrast Promo Package |  | AddPart:79035,MFA34100,MFV30010,MMK30105,MMK30105,MRP00042,MRP20102,MRP26202
999-368 | Ts2 Tissue Culture ADL Phase Conrast Promo Package |  | AddPart:79035,MFA34100,MFV30010,MMK30105,MMK30105,MRP00042,MRP40102,MRP46202
999-369 | Ts2-FL 470nm Phase Promo Package - Add Optics |  | AddPart:79035,96242,MBF84735,MFA34200,MFV30010,MFV51000,MMK30105,MMK30105
999-370 | Ts2-FL 470nm 10x Phase Promo Package |  | AddPart:79035,96242,MBF84735,MFA34200,MFV30010,MFV51000,MMK30105,MMK30105,MRP40102
999-371 | Ts2-FL 470nm 10x & 20x Phase Promo Package |  | AddPart:79035,96242,MBF84735,MFA34200,MFV30010,MFV51000,MMK30105,MMK30105,MRP40102,MRP46202
899-438 | Nikon Ti2 with AX Lite |  | AddPart:MHA52000,MHE60320,77013809,77013809,MHE60200,MHE60220,MHE60220,98303,MHV55010,MXA22194,MHF49200,MHF49210,MQF52057,MHF49000,MHF46800,MHS50000,77019803,77019820,79035,79035,79035,79035,79035,79035,MEE55705,MED54555,MEA54000,MEF55037,MEB55835,MEB52340,MAK10110,MAK10110,MEE59925,MEL51005,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084,97050,97050,MBF83000,MQF52057,MXA22196,96398,MXA22142,MXA22143,MEC51210,MEC57110,MEC55016,MEP59320,77050322
899-439 | Nikon Ti2 with AX Promotion |  | AddPart:MHA52000,MHE60320,77013809,77013809,98701,98303,MHV55010,MXA22194,MHF49200,MHF49210,MQF52057,MHF49000,MHF46800,MHS50000,MQS43130,MQS43200,77010604,77019803,77019820,79035,79035,79035,79035,79035,MEE55855,MHE50350,MED54555,MEA54000,MEF55037,MEB55835,MEB52520,MAK10110,MAK10110,MEP59394,MEE59925,MEL51920,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,MXA22205,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084,MEC56120,MEF55706,MEC59140,MEC59150,97050,77050313,MRD70040,MRD70170,MRD70270,MRD71670,77060136,MQS41920,96370,96372,96375,96376
899-440 | AXR Upgrade Promotion |  | AddPart:MHA52100,MHE60310,MHE60220,MHE60220,MHE60210,MHE60210,MHE60000,99343,98312,77013809,77013809,MXA22194,MHF49200,MHF49210,MQF52057,MHF49000,MHF46800,77011682,MHS50000,77019803,77019820,79035,79035,79035,79035,79035
899-441 | Nikon Ji with AX Promo |  | AddPart:MHA52000,MHE60320,77013809,77013809,98701,98303,MHV55010,MXA22194,MHF49200,MHF49210,MHF49000,MHF46800,MHS50000,MQS43130,MQS43200,77010604,77019803,77019820,79035,79035,79035,79035,79035,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,MQF52057,97050,77050313,MFA92000,MFF92000,MFD92010,MFC92000,MFD92020,MBF83200,96378,MQS44000,MRD70040,MRD70170,MRD70270,MRD71670
899-442 | Nikon Ji with AXR Promo |  | AddPart:MHA52100,MHE60310,77013809,77013809,98701,98303,MHV55010,MXA22194,MHF49200,MHF49210,MHF49000,MHF46800,MHS50000,MQS43130,MQS43200,77010604,77019803,77019820,79035,79035,79035,79035,79035,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,MQF52057,97050,77050313,MFA92000,MFF92000,MFD92010,MFC92000,MFD92020,MBF83200,96378,MQS44000,MRD70040,MRD70170,MRD70270,MRD71670
899-443 | AXR + NSPARC Ti2 System Promotion |  | AddPart:MHE62000,MHE63000,MHE63010,MHE63020,MHE63030,MHA52100,MXA22211,MXA22194,MHF49200,MHF49210,MHF49030,MHF46810,MQF52057,MHS50000,77019803,77019820,79035,79035,79035,79035,79035,77013809,77013809,MQS43130,MEE55705,MHV55010,MED54555,MEA54000,MEF55037,MEB55835,MEB52520,MAK10110,MAK10110,MEP59394,MEE59925,MEL51920,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,MXA22205,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084,MEC56120,MEF55706,MEC59140,MEC59150,97050,77050313,MRD70040,MRD70170,MRD70270,MRD71670,77060136,MQS41920,77011682,96370,96372,96375,96376
899-444 | AX + NSPARC Ti2 System Promotion |  | AddPart:MHE62000,MHE63000,MHE63010,MHE63020,MHE63030,MHA52000,MXA22211,MXA22194,MHF49200,MHF49210,MQF52057,MHF49030,MHF46810,MHS50000,77019803,77019820,79035,79035,79035,79035,79035,77013809,77013809,MQS43130,MEE55705,MHV55010,MED54555,MEA54000,MEF55037,MEB55835,MEB52520,MAK10110,MAK10110,MEP59394,MEE59925,MEL51920,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,77013066,77013066,76998,76998,76998,76998,76998,76998,76998,76998,76998,76998,92084,MEC56120,MEF55706,MEC59140,MEC59150,97050,77050313,MRD70040,MRD70170,MRD70270,MRD71670,77060136,MQS41920,96370,96372,96375,96376,MXA22205
899-445 | AXR NSPARC Add-on Promotion |  | AddPart:MHE62000,MHE63000,MHE63010,MHE63020,MHE63030,MHA52100,MXA22194,MHF49200,MHF49210,MQF52057,MHF49030,MHF46810,MHS50000,77019803,77019820,79035,79035,79035,79035,79035,77013809,77013809
899-446 | NSPARC Upgrade for AX |  | AddPart:MHE62000,MHE63000,MHE63010,MHE63020,MHE63030,77012928,C3700513
899-447 | Crest V3 DeepSIM-Ready Promo |  | AddPart:77060140,77081013,77081014,MXA22158,77059524,MED54555,MQS31000,MQS42560,77010640,MQS41930,MQS41220,77015917,77015914,77015916,77015913,77015839,77015565,77081023,MEA54000,MEF55037,97050,79035,79035,79035,79035,79035,MEB55835,MEB52340,MAK10110,MAK10110,MEP59395,MEE59925,MEE55705,MEL51005,MEL56200,MEV51030,MXA22147,MEE54860,MEE54741,MXA22205,77060136,MQS41920,96370,96372,96375,96376,MEC56120,MEF55706,MEC59140,MEC59150,92084,77023034,MQS42780,77059528,77059529
899-448 | AXR-MP Ready Upgrade Promotion |  | AddPart:MHA52300,MHE60310,MHE60220,MHE60220,MHE60210,MHE60210,MHF49000,MHF46800,MHE60000,99343,98312,77013809,77013809,MXA22194,MHF49200,MHF49210,MQF52057,MHS50000,77019803,77019820,79035,79035,79035,79035,79035
899-449 | Nikon Fi3 Color Camera Promotion |  | AddPart:MQA18000,MQD42075,MQF52057,79035,77013245
899-450 | Nikon E-TIRF Upgrade Promotion |  | AddPart:MEE59220,MXA22147,MEE54840,77023046,MEE59178,77097070,79035,MEE58017,97338,97062,MEB55855
899-451 | AXR-MP 920nm Promotion |  | AddPart:77049818,MHA52300,MHE60560,79035,79035,79035,79035,79035,99194,MHE61020,MHV55560,MHV55550,MBD95000,MBD92300,MBF95600,MBF95500,MBF95200,77011694,MRD77226,MBE94300,MBE94150,77060136,MQS41920,96372,96375,MXA22205,MBB94511,MHV55720,MAK10110,MAK10110,MBP70110,MQS41100,99122,MHE60500,MHE60420,MHF46800,MXA22194,MHF49200,MQF52057,MHF49210,MHV53040,77011695,77012801,MHS50000,MQS41100,77019821,77019803,77019793,77019802
899-452 | AXR-MP Dual Line Promotion |  | AddPart:77049814,MHA52300,MHE60560,79035,99709,MHE61000,MHV55530,MHE61010,MBD95000,MBD92300,MBF95600,MBF95500,MBF95200,77011694,MBE94300,MBE94150,77060136,MQS41920,96372,96375,MXA22205,MBB94511,MHV55720,MAK10110,MAK10110,MBP70110,MRD77226,MQS41100,99321,99332,99333,MHE60500,MHE60430,MHE60530,MHE60530,MHF46800,MXA22194,MHF49200,MQF52057,MHF49210,MHV53040,77011695,77012801,MHS50000,MQS41100,77019821,77019803,77019793,77019802
899-453 | Crest DeepSIM Upgrade Promotion |  | AddPart:77059520,C3700530,77019808
899-454 | Super STORM Ti2 Upgrade Promo |  | AddPart:77019821,77019803,97050,79035,79035,79035,79035,79035,79035,79035,79035,79035,79035,77023034,77097051,77026047,MEE54741,MEE59220,MEE54330,MEE54850,MEE54880,MEE58017,MED54301,91989,91995,91997,MRD01995,97327,MEE59178,MEE59117,97062,MEB55855,MQS31000,MQS42560,MQS41930,MQS42780,MQS42930,MQS42650,MEE54741,77023046

## PCs
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
77019680 | DEMO 64BIT PC Z420, 16GB RAM 1TB WIN 7 |  | DEMO 64 bit Imaging workstation supplied with Intel Xeon 3.6GHz quadcore processor, 16GB DDR3 RAM, 1TB hard drive (mirrored RAID for dataintegrity), and 1GB video card. Optical DVD-RW, gigabit LAN adapter,x5 USB2.0 and x4 USB3.0 ports, x1 IEEE1394a port, USB keyboard and lasermouse, and recover software are included. Windows 7 64bit Professionaledition installed. NOTE: ONLY AVAILABILITY IS FROM CONSIGNMENT PURCHASEOR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019684 | DEMO 64BIT Z420 32GB 1TBOS/3TB D DR WIN7 |  | DEMO 64 bit Imaging workstation supplied with Intel Xeon 3.7GHz quadcore processor, 32GB DDR3 RAM, dual 7200RPM SATA drives: a 1TB OS and3TB data drive, and 1GB video card. Optical DVD-RW, gigabit LANadapter,x5 USB2.0 and x4 USB3.0 ports, keyboard and laser mouse, andrecover software are included. Windows 7 64bit Professional editioninstalled. Kit includes display port to DVI-D adapter. NOTE: ONLYAVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019697 | DEMO 64BIT Z440 32GB/4TB DATA DR WIN7 |  | DEMO 64bit Imaging workstation with Xeon 3.7GHz 4 Core CPU, 32GB DDR4RAM, 4TB (2x2TB striped RAID) Hard drive, 2GB Video, optical DVDRW, x2LAN ethernet adapters, x8 USB 3 and x2 USB 2 ports, keyboard and laseroptical mouse, and recovery software included. Windows 7 64bitprofessional edition. NOTE: ONLY AVAILABILITY IS FROM CONSIGNMENTPURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019698 | DEMO 64BIT Z440 64/256GB SS DR/4TB DR W7 |  | DEMO 64bit Imaging workstation with Xeon 3.0GHz 8 Core CPU, 64GB DDR4RAM, a 256GB SSD OS hard drive, 4TB (2x2TB striped RAID) data storagehard drive, 4GB Video, optical DVDRW, x2 LAN ethernet adapters, x8 USB 3and x2 USB 2 ports, keyboard and laser optical mouse, and recoverysoftware included. Windows 7 64bit professional edition. NOTE: ONLYAVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
99903 | DEMO CONFOCAL Z400 24GB RAM WIN7 PKG |  | "DEMO 64bit Confocal PC Workstation supplied with Intel Xeon 3.20 GHz (6core, 12MB cache) processor, 24GB DDR3 RAM, USB keyboard and lasermouse, and 3D mouse. 1GB display adapter (supports two monitors), SATABluRay Writer, 2x300GB SATA drives (10K RPM driv" | Monitor:1:4:H | PCIEport:0:5:H
77019781 | DEMO HP Z43 42.5-INCH 4K UHD DISPLAY |  | DEMO HP Z43 42.5-inch 4K UHD Display. NOTE: ONLY AVAILABILITY IS FROMCONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019706 | DEMO HP Z440 CONFOCAL WORKSTATION FOR SR |  | DEMO 64bit Imaging workstation optimized for confocal and deconvolutionwith Xeon 3.0GHz 8 Core CPU, 64GB DDR4 RAM, a256GB SSD OS hard drive,4TB (2x2TB striped RAID) data storage harddrive, 6GB Video for high endGPU processing, optical DVDRW, x2 LAN ethernet adapters, x8 USB 3 andx2USB 2 ports, keyboard and laser optical mouse, and recoverysoftwareincluded. Windows 7 64bit professional edition. NOTE: ONLYAVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019657 | DEMO IMAGING WORKSTATION, Z400-8GB |  | DEMO Imaging Workstation supplied with HP Z400. Installed with WindowsVista 64 Bit Business Edition, Intel Xeon W3505 2.53 4MB/1066 DC CPU,8GB 2x4GB DDR3-1333 ECC RAM, NVIDIA Quadro FX580 512MB Graphics card,Gigabit LAN adapter, USB keyboard and optical scroll mouse, HP 16XDVDwriter uperMulti SATA 1st, 2 HP 500GB SATA Drives, RAID 0 Striped.DISCONTINUED; - AVAILABLE ONLY FROM DEMO INVENTORY OR CONSIGNMENT;77019664 IS REPLACEMENT | Monitor:1:4:H | PCIEport:0:5:H
77019716 | DEMO Z440 ENHANCED IMAGING WORKSTATION |  | DEMO Enhanced Imaging Workstation - HP Z440 featuring Intel Xeon 3.5 GHzprocessors (4 cores), 32GB DDR4 2400 Mhz RAM, GTX 1080 8GB GDDR5X GPU,4TB hard drive (2GBx2 Raid 0), 2x Gb LAN ethernet, 8x USB3 and 2x USB2ports, keyboard and optical mouse. Windows 10 64-bit Professional.LIMITED AVAILABILITY FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY | Monitor:1:4:H | PCIEport:0:5:H
77019725 | DEMOZ440 ENHANCED IMAGING WORKSTATION V3 |  | DEMO Enhanced Imaging Workstation - HP Z440 featuring Intel Xeon 3.5GHz processors (4 cores), 32GB DDR4 2400 Mhz RAM, 64-bit Windows 10Professional on 512GB SSD, 4TB data hard drive array (2GBx2 Raid 0), 8GBGDDR5 Professional Rendering GPU, 2x Gb LAN ethernet, 8x USB3 and 2xUSB2 ports, keyboard and optical mouse. LIMITED AVAILABILITY FROMCONSIGNMENT PURCHASE OR DEMO INVENTORY | Monitor:1:4:H | PCIEport:0:5:H
77019730 | DEMOZ440 ENHANCED IMAGING WORKSTATION V4 |  | DEMO Enhanced Imaging Workstation - HP Z440 featuring Intel Xeon 3.7GHz processors (4 cores), 32GB DDR4 2400 Mhz RAM, 64-bit Windows 10Professional on 512GB SSD, 4TB data hard drive array (2GBx2 Raid 0), 8GBGDDR5 Professional Rendering GPU, 2x Gb LAN ethernet, 8x USB3 and 2xUSB2 ports, keyboard and optical mouse. LIMITED AVAILABILITY FROMCONSIGNMENT PURCHASE OR DEMO INVENTORY | Monitor:1:4:H | PCIEport:0:5:H
77019746 | DEMO Z4 ADVANC BIOSYSTEMS WORKSTATION V2 |  | Advanced Imaging Workstation - HP Z4 featuring Intel Xeon 3.7 GHzprocessor (8 cores), 64GB DDR4 2666 Mhz RAM, NVIDIA Professional 8GBGDDR5X GPU, 512GB M.2 Turbo SSD O/S drive, 1TB M.2 Turbo SSD data drive,12TB data storage, 2x Gb LAN ethernet, premium keyboard and opticalmouse. Windows 10 64-bit Professional.LIMITED AVAILABILITY FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY ONLY | Monitor:1:4:H | PCIEport:0:5:H
77019757 | DEMO Z4 ADVANCED BIOSYS WORKSTATION V3 |  | DEMO Advanced Biosystems Z4 Workstation - HP Z4 featuring Intel Xeon 3.7GHz processor (8 cores), 128GB DDR4 2666 Mhz RAM, NVIDIA RTX5000 16GB, 512GB M.2 Turbo SSD O/S drive, 1TB M.2 Turbo SSD data drive, 12TB datastorage, 2x Gb LAN ethernet/10GbE Dual Port NIC, premium keyboard andoptical mouse. Windows 10 64-bit Professional. NOTE: ONLY AVAILABILITYIS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019767 | DEMO Z4 ADVANCED BIOSYS WORKSTATION V4 |  | DEMO Advanced Biosystems Z4 Workstation - HP Z4 featuring Intel Xeon3.9GHz processor (8 cores), 128GB DDR4 2933 Mhz RAM, NVIDIA RTX500016GB, 512GB M.2 Turbo SSD O/S drive, 1TB M.2 Turbo SSD data drive, 12TBdata storage, 2x Gb LAN ethernet/10GbE Dual Port NIC, premium keyboardand optical mouse. Windows 10 64-bit Professional. NOTE: ONLYAVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019756 | DEMO Z4 ADVANCED IMAGING WORKSTATION V3 |  | DEMO Advanced Imaging Workstation - HP Z4 featuring Intel Xeon 3.7 GHzprocessor (8 cores), 64GB DDR4 2666 Mhz RAM, NVIDIA RTX4000 8GB, 512GBM.2 Turbo SSD O/S drive, 1TB M.2 Turbo SSD data drive, 12TB datastorage, 2x Gb LAN ethernet, premium keyboard and optical mouse. Windows10 64-bit Professional. NOTE: ONLY AVAILABILITY IS FROM CONSIGNMENTPURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019763 | DEMO Z4 ADVANCED IMAGING WORKSTATION V4 |  | DEMO Z4 ADVANCED IMAGING WORKSTATION V4 - NOTE: ONLY AVAILABILITY ISFROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019748 | DEMO Z4 ENHANCED IMAGING WORKSTATION V2. |  | Enhanced Imaging Workstation - HP Z4 featuring Intel Xeon 3.6 GHzprocessor (4 cores), 32GB DDR4 2666 Mhz RAM, P4000 8GB GDDR5X GPU, 512GBTurbo M.2 OS drive, 2x2TB HDD storage, 2x Gb LAN ethernet, 10x USB3.1ports, keyboard and optical mouse. Windows 10 64-bit Professional.LIMITED AVAILABILITY FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY ONLY | Monitor:1:4:H | PCIEport:0:5:H
77019755 | DEMO Z4 ENHANCED IMAGING WORKSTATION V3 |  | DEMO Enhanced Imaging Workstation - HP Z4 featuring Intel Xeon 3.6 GHzprocessor (4 cores), 32GB DDR4 2666 Mhz RAM, NVIDIA RTX 4000 8GB, 512GBTurbo M.2 OS drive, 2x2TB HDD storage, 2x Gb LAN ethernet, 10x USB3.1ports, keyboard and optical mouse. Windows 10 64-bit Professional.NOTE: ONLY AVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019771 | DEMO Z4 ENHANCED IMAGING WORKSTATION V4 |  | DEMO Z4 ENHANCED IMAGING WORKSTATION V4; HP Z4G4T XW2223 32GB/2TB NOTE:ONLY AVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019776 | DEMO Z4 ENHANCED IMAGING WORKSTATION V5 |  | DEMO Enhanced Imaging Workstation - Minimum specification based on HP Z4featuring Xeon 3.6GHz CPU, 32GB RAM, NVIDIA RTX A4000 GPU, 512GB SSD OSdrive, 2x2TB HDD storage, 2x 1GbE LAN ports, 10x USB 3.1ports Type A,DVD-RW, keyboard and optical mouse. Windows 10 64-bit Professional. NOTE: ONLY AVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMOINVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019785 | DEMO Z4 ENHANCED IMAGING WORKSTATION V6 |  | DEMO Enhanced Imaging Workstation - Minimum specification based on HP Z4featuring Xeon 3.6GHz CPU, 32GB RAM, NVIDIA RTX A4000 GPU, 512GB SSD OSdrive, 2x2TB HDD storage, 2x 1GbE LAN ports, 10x USB 3.1ports Type A,DVD-RW, keyboard and optical mouse. Windows 10 64-bit Professional.NOTE: ONLY AVAILABILITY IS FROM CONSIGNMENT PURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019790 | DEMO Z4 ENHANCED IMAGING WORKSTATION V7 |  | DEMO HP Z4 with Xeon 4.2G processor-6 core, 128GB RAM (upgradeable to256GB),NVIDIA RTX A4000 16GB GPU, 1TB M.2 DDS OS Drive, 12TB SATAdatadrive, 2x 10GB Ethernet ports, 8x USB ports Type A, 2x 20Gbps USB-C,DVD-RW, premium keyboard and optical mouse. Optiona Wifi card.Windows 1064-bit Professional. NOTE: ONLY AVAILABILITY IS FROM CONSIGNMENTPURCHASE OR DEMO INVENTORY. | Monitor:1:4:H | PCIEport:0:5:H
77019720 | Z440 ENHANCED IMAGING WORKSTATION V2 |  | Enhanced Imaging Workstation - HP Z440 featuring Intel Xeon 3.5 GHzprocessors (4 cores), 32GB DDR4 2400 Mhz RAM, 4TB hard drive (2GBx2 Raid0), 2x Gb LAN ethernet, 8x USB3 and 2x USB2 ports, keyboard and opticalmouse. Windows 10 64-bit Professional. Includes 3 years warranty forparts, labor, and service. Includes GTX1080 Video Card. | Monitor:1:4:H | PCIEport:0:5:H
77019810 | DOC WORKSTATIONS Z2G5 MINI V4 |  | HP Z2 G9 with i7-13700K 3.40G 16 core CPU, 16GB RAM, 512GB M.2 SSD,NVIDIA T1000 4GB GPU with 4 display ports, 4x USB-A, 3x USB-C20Gbps,1GbE LAN, Internal wifi, slim keyboard, mouse. Windows 11 Pro, HP3/3/3warranty.Does not take PCIe cards. | Monitor:1:4:H | PCIEport:0:5:H
77019733 | DEMO Z4 ENHANCED IMAGING WORKSTATION |  |  | Monitor:1:4:H | PCIEport:0:5:H
77019770 | DEMO Z4 ADVANCED IMAGING WORKSTATION V4 |  |  | Monitor:1:4:H | PCIEport:0:5:H
77019805 | DOC WORKSTATION Z2 TOWER V3 |  | HP Z2 G9 Tower with i5-13600K 2.70G 14 core CPU, 64GB RAM (upgradeableto 128), NVIDIA RTX A4000 16GB GPU, 512GB M.2 SSD, 12TB SATA data drive,10x USB 3.1 ports Type A, 1x USB-C 20Gbps, DVD-RW, Internal Wifi,premium keyboard and optical mouse. Windows 11 Pro, HP 3/3/3warranty. Not recommended for use as an imaging level workstation. | Monitor:1:4:H | PCIEport:0:5:H
77019820 | Z4 ENHANCED IMAGING WORKSTATION V9 |  | HP Z4 with Xeon 4.2G processor-6 core, 128GB RAM (upgradeableto256GB),NVIDIA RTX A4000 16GB GPU, 1TB M.2 DDS OS Drive, 12TBSATAdatadrive, 2x 10GB Ethernet ports, 8x USB ports Type A, 2x 20GbpsUSB-C,DVD-RW, premium keyboard and optical mouse. Optiona Wificard.Windows 11 64-bit Professional. HP 3/3/3 warranty. | Monitor:1:4:H | PCIEport:0:5:H
77019806 | Z4 ENHANCED IMAGING WORKSTATION V8 |  | HP Z4 with Xeon 4.2G processor-6 core, 128GB RAM (upgradeable to256GB),NVIDIA RTX A4000 16GB GPU, 1TB M.2 DDS OS Drive, 12TB SATAdatadrive, 2x 10GB Ethernet ports, 8x USB ports Type A, 2x 20Gbps USB-C,DVD-RW, premium keyboard and optical mouse. Optiona Wifi card.Windows 11 64-bit Professional. HP 3/3/3 warranty. | Monitor:1:4:H | PCIEport:0:5:H
77019821 | Z4 ADVANCED IMAGING WORKSTATION V8 |  | HP Z4 with Xeon 4.6G processor-20 core, 256GB RAM (upgradeableto512GB),NVIDIA RTX A4000 16GB GPU, 1TB M.2 SSD OS drive, 12TBSATA datadrive,2x 10GB Ethernet ports, 8x USB Type A, 2x20Gbps USB-C, DVD-RW,premiumkeyboard and optical mouse. Optional Wificard. Windows 11 64-bitProfessional. HP 3/3/3 warranty. | Monitor:1:4:H | PCIEport:0:5:H
77019807 | Z4 ADVANCED IMAGING WORKSTATION V7 |  | HP Z4 with Xeon 4.6G processor-20 core, 256GB RAM (upgradeable to512GB),NVIDIA RTX A4000 16GB GPU, 1TB M.2 SSD OS drive, 12TBSATA data drive,2x 10GB Ethernet ports, 8x USB Type A, 2x20Gbps USB-C, DVD-RW, premiumkeyboard and optical mouse. Optional Wificard. Windows 11 64-bitProfessional. HP 3/3/3 warranty. | Monitor:1:4:H | PCIEport:0:5:H
77019808 | Z8 FURY WORKSTATION ENHANCED V2 |  | HP Z8 FURY with Intel Xeon W5-3435X 4.50G-16 core, 512GB RAM(upgradeableto 1TB), NVIDIA RTX 5000 ADA 32GB GPU, 2TB M.2 SSD, 12TB SATAdata drive, 2x10GB ethernet, 8x USB Type A, 20Gbps USB-C, DVD-RW,premium keyboard andoptical mouse. Optional Wifi Card. Windows 11 64-bitProfessional. HP3/3/3 warranty. | Monitor:1:4:H | PCIEport:0:5:H
77019792 | Z8 FURY WORKSTATION ADVANCED |  | HP Z8 FURY with Xeon 4.6G processor-56 Core, 512GB RAM (upgradeable to 1TB), NVIDIA RTX 6000 Ada 48GB GPU, 2x 2TB M.2 SSDs, 2x 12TB sata data drives, 2x 10GB ethernet, 8x USB Type A, 2x 20Gbps USB-C, DVD-RW, premium keyboard and optical mouse. Optional Wifi Card. Windows 11 64-bit | Monitor:1:4:H | PCIEport:0:5:H
77019764 | DOCUMENTATION WORKSTATION HP Z2G5/T1000 |  |  | Monitor:1:4:H | PCIEport:0:5:H
77019791 | HP Z4 with Xeon 4.6G processor-20 core, 256GB RAM (upgradeable to512GB),NVIDIA RTX A4000 16GB GPU, 1TB |  |  | Monitor:1:4:H | PCIEport:0:5:H
77019778 | Z4 ADVANCED BIOSYSTEMS WORKSTATION V5 |  |  | Monitor:1:4:H | PCIEport:0:5:H
77019777 | Advanced Imaging Workstation # Minimum specification based on HP Z4featuring Xeon 3.8GHz CPU, 64GB RAM |  |  | Monitor:1:4:H | PCIEport:0:5:H
77019811 | 24" IPS PRO 7 WUXGA MONITOR |  |  | Monitor:1:1:P
77019813 | 27" QHD MONITOR, 727PQ |  |  | Monitor:1:1:P
77019812 | 27" IPS PRO THUNDERBOLT 4K DISPLAY |  |  | Monitor:1:1:P
77019809 | 39.7" ULTRA WIDE CURVE DISPLAY, 5K2K |  |  | Monitor:1:1:P
77019803 | THUNDERBOLT 31.5 INCH 4K DISPLAY |  |  | Monitor:1:1:P
77019731 | 37.5-inch IPS, LED Backlit Monitor |  |  | Monitor:1:1:P
77019765 | DEMO 24" LED IPS MONITOR, Z24N G3 |  |  | Monitor:1:1:P
77019738 | 31.5" Z32 4k UHD Display. |  |  | Monitor:1:1:P
77019768 | DEMO 27-inch IPS, LED Backlit Monitor |  |  | Monitor:1:1:P
77019769 | DEMO 27" Z27K G3 4K UHD DISPLAY |  |  | Monitor:1:1:P
77019732 | 27" LCD MONITOR, Z27N G2 |  |  | Monitor:1:1:P

## InterlockandTriggering
Product Code | Description | Tags | Notes | slot1 | slot2 | slot3 | slot4 | slot5 | slot6 | slot7 | slot8 | slot9 | slot10 | slot11 | slot12 | slot13 | slot14 | slot15 | slot16
77023046 | INTERLOCK JUNCTION BOX | Interlock | 2 Phone + 2 Round for lasers, 1 Phone or Round for Mic, 1 Phone or Round for LSC | InterlockRound:0:3:H | InterlockPhone:0:3:H | InterlockPhone¦InterlockRound:0:1:H
MHF49200 | LU-INTL INTERLOCK BOX | LaserSafety |  | AXInterlockBox:1:1:P | InterlockPlug:1:7:H | LaserIndicator:1:1:H | InterlockRound:1:1:H
77012519 | AX LUNF INTERLOCK CABLE | Interlock | LUNF to Plug | InterlockLUNF:0:1:P | InterlockPlug:0:1:P
77012524 | AX AUXILIARY INTERLOCK CABLE | Interlock | Phone to Plug | InterlockPhone:0:1:P | InterlockPlug:0:1:P
MXA22094 | A1 TI INTERLOCK CABLE | Interlock | Phone to Phone | InterlockPhone:0:1:P | InterlockPhone:0:1:P
MXA22158 | S-TI2-LU INTERLOCK CABLE | Interlock | Round to Phone | InterlockRound:0:1:P | InterlockPhone:0:1:P
MXA22194 | LU-INTL-C INTERLOCK CABLE | Interlock | Round to Round | InterlockRound:0:1:P | InterlockRound:0:1:P
MEE59178 | TI2-LU-LSC LASER SAFETY COVER | Interlock,LaserSafety,LaserSafetyCover | Round | InterlockRound:0:1:P
MEE59177 | TI-LA-LASER SAFETY COVER | Interlock,LaserSafety,LaserSafetyCover | Phone | InterlockPhone:0:1:P
MHF49210 | LU-IND LASER INDICATOR | Interlock | Plug | LaserIndicator:1:1:P | ACAdapter:1:1:H
77012509 | LASER INTERLOCK CABLE F/LUN-F | Interlock | Phone to LUNF | InterlockLUNF:0:1:P | InterlockPhone:0:1:P
JBAM-LUNFGenderChanger | LUNF Cable Gender Changer | Interlock | Host LUNF to Host Phone | InterlockPhone:0:1:H | InterlockPhone:0:1:H